Computational Design of Biodegradable Metallic Implants: Modeling, Optimization, and Clinical Translation

Henaka Ariyarathna​a,b, Aijun Huang​a,b,c, Yunlong Tang​a,b,c*, Yuman Zhu​a,c*

​aMonash Centre for Additive Manufacturing, Monash University, Wellington Rd, Clayton, VIC 3800, Australia

​bDepartment of Materials Science and Engineering, Monash University, Wellington Rd, Clayton, VIC 3800, Australia

​cDepartment of Mechanical and Aerospace Engineering, Monash University, Wellington Rd, Clayton, VIC 3800, Australia

Abstract

Zinc has emerged as a promising biodegradable metal for temporary implant applications. Despite extensive experimental studies, predictive computational frameworks for Zn biodegradation under physiological conditions remain limited. In this study, a three-dimensional mechanistic reaction-diffusion framework is developed by coupling oxygen reduction reaction (ORR) kinetics, mass transport, corrosion-product evolution, and moving-interface tracking. Initial sensitivity analysis showed that dissolved oxygen availability as the primary control on Zn biodegradation. The model was calibrated against literature-reported immersion mass-loss data and predicted 28-day immersion experiments of pure Zn in Hanks' Balanced Salt Solution, yielding a root-mean-square error of 0.0159%. Application of the calibrated model to a six-crown Zn stent, without parameter refitting, produced degradation trends consistent with reported in vivo observations and demonstrated that geometry-dependent material loss emerges naturally from local variations in oxygen transport and corrosion-product evolution. Overall, the proposed framework establishes a predictive foundation for the design and evaluation of biodegradable Zn implants.

Keywords: Biodegradable implant; Zinc; Corrosion modeling; Oxygen reduction kinetics; Ionic transport

# Introduction

## Clinical motivation and limitations of permanent implants

Writing prompt: Establish the clinical need for temporary metallic support, including avoidance of removal surgery, long-term stress shielding, and persistent-device complications.

[Add review text and citations here]

Modern medicine increasingly relies on metallic implants to restore physiological function and provide structural support in orthopedic, cardiovascular, and reconstructive applications. Metals are particularly suited to these applications because their high strength, fracture toughness, fatigue resistance, and load-bearing capacity generally exceed those of polymers and ceramics [1]. Conventional implant materials, including stainless steels, titanium alloys, and cobalt–chromium alloys, were developed to resist corrosion and maintain long-term mechanical stability in physiological environments [2–5]. Although this durability has enabled their widespread clinical use, permanent retention can be undesirable when an implant is required only during a limited period of tissue healing. Long-term implant presence may contribute to stress shielding, chronic local reactions associated with wear or metal-ion release, interference with tissue growth and remodeling, and vascular complications such as thrombosis or restenosis. Moreover, temporary fixation devices may require surgical removal after healing, increasing the risks of anesthesia, infection, additional tissue damage, prolonged recovery, and healthcare costs [6–10]. Biodegradable metallic implants have therefore emerged as an alternative strategy to provide mechanical support during the critical healing period and then degrade into products the body can tolerate, metabolize, or clear [11–15]. Their gradual degradation may eliminate the need for removal surgery while progressively transferring mechanical load from the implant to the recovering tissue. Magnesium (Mg), iron (Fe), and zinc (Zn), together with their alloys, are the principal metallic systems investigated for biodegradable implant applications because they combine clinically relevant mechanical properties with the capacity to degrade under physiological conditions [16–20]. Researchers are exploring these materials for orthopedic fixation devices, vascular scaffolds, porous bone implants, and other temporary biomedical devices. However, their degradation behavior and functional performance depend strongly on material composition, microstructure, surface condition, implant geometry, manufacturing route, mechanical loading, and the surrounding biological environment. Developing biodegradable metallic implants therefore requires an integrated design approach that provides sufficient mechanical support during healing while controlling the rate, morphology, and biological consequences of degradation.

Unlike permanent devices, biodegradable metallic implants are dynamic systems whose geometry, surface condition, mechanical capacity, and biological interactions evolve continuously after implantation. Coupled electrochemical reactions, mass transport, corrosion-product formation, mechanical loading, and local tissue responses govern their degradation, and material composition, microstructure, manufacturing history, and implant geometry further influence it [21–23]. Consequently, degradation measured using simplified laboratory specimens may not represent the behavior of functional screws, plates, scaffolds, or stents in vivo [21,24]. Importantly, average mass loss alone cannot define implant performance. Uniform surface recession and localized corrosion can produce similar total mass loss but markedly different residual strength, particularly when degradation occurs at critical features such as screw threads, plate holes, scaffold junctions, or stent crowns [25–28]. A successful implant must therefore retain adequate mechanical function during the critical healing period while gradually transferring load to the recovering tissue and degrading without harmful accumulation of corrosion products. Biodegradable implants should therefore be designed around their complete degradation–mechanical performance trajectory, rather than their initial properties or nominal corrosion rate alone. This time-dependent behavior requires an integrated design framework that considers degradation, mechanics, transport, manufacturing, and tissue healing as interacting components of a single material–device system [22,24,29].

The design of biodegradable metallic implants is governed by four interdependent mismatches. First, implant degradation must be synchronized with tissue healing: premature degradation may cause loss of fixation or structural support, whereas prolonged persistence may negate the advantages of a temporary device [30,31]. Second, the progressive reduction in implant strength and stiffness must be balanced by the recovery of the surrounding tissue to ensure gradual load transfer without excessive deformation or sudden instability [23,32,33]. Third, ion and corrosion-product release must remain compatible with transport, biological tolerance, metabolism, and clearance; otherwise, local accumulation may occur even when the average degradation rate appears acceptable [31,34,35]. Finally, computational models must balance physical fidelity with experimental and computational feasibility. Although detailed multiphysics models can represent electrochemical reactions, mass transport, film evolution, mechanics, and tissue response, their parameter requirements and computational cost can limit calibration, uncertainty analysis, and design optimization [36–38]. Effective implant design therefore requires models that capture the mechanisms controlling clinically relevant outcomes, remain identifiable from available experiments, and are efficient enough for device-scale simulation.

This review critically examines the computational design of biodegradable metallic implants across Mg-, Fe-, and Zn-based systems, emphasizing the relationships among material properties, degradation behavior, evolving mechanical performance, implant geometry, manufacturing, and tissue healing. It covers empirical, phenomenological, stochastic, and mechanistic degradation models; degradation-coupled structural and mechanobiological simulations; porous and topology-optimized architectures; additive-manufacturing-aware design; surrogate modeling; machine learning; uncertainty quantification; and model verification and validation [32–38]. Rather than organizing these developments into separate material- or method-specific topics, the review introduces a unified, lifetime-aware computational design framework that evaluates implants over their full functional period. Within this framework, material selection and implant geometry determine the initial response; degradation models predict the evolution of morphology and material loss; mechanical models quantify the resulting decline in device function; and tissue-healing models describe the progressive recovery of biological load-bearing capacity. Optimization then identifies designs that maintain application-specific performance throughout healing while satisfying degradation, biological, manufacturing, and safety constraints. The principal novelty of this review is integrating previously fragmented computational approaches into a common pathway linking material-scale degradation mechanisms to device-scale performance and clinical translation. This review pays particular attention to model transferability among Mg, Fe, and Zn systems; the identification and calibration of degradation parameters; the limitations of extrapolating from coupons to functional devices; and the experimental and uncertainty evidence needed for computational predictions to support clinically relevant design decisions.

This review follows the complete computational design pathway for biodegradable metallic implants. It first defines application-specific clinical requirements and the time-dependent relationship between implant degradation, mechanical support, and tissue healing. The review then compares Mg-, Fe-, and Zn-based systems from a computational perspective and examines empirical, phenomenological, stochastic, and mechanistic approaches to modeling degradation. Subsequent sections address the coupling of degradation with structural mechanics, implant–tissue interactions, and evolving functional performance, followed by computational methods for geometry optimization, porous and graded architecture design, and additive-manufacturing-aware development. The review then evaluates the roles of sensitivity analysis, surrogate modeling, machine learning, uncertainty quantification, and multi-objective optimization. Finally, the review considers model calibration, verification and validation, experimental-to-device translation, regulatory credibility, and the remaining barriers to clinical implementation. This sequence connects fundamental degradation physics with implant-level design decisions and provides a structured pathway toward predictive, lifetime-aware, and clinically translatable computational design of biodegradable metallic implants.

# Clinical design requirements

## Application classes

Table 1 summarizes the application-specific mechanical, temporal and degradation requirements discussed in the following subsections; requirements common to all devices (stiffness compatibility with bone, 4–6 weeks of full support, and resorption after functional healing) are stated once in the overview of orthopedic fixation devices and not repeated for each device.

Table 1. Application-specific requirements for biodegradable metallic implants (YS: yield strength; UTS: ultimate tensile strength; SS: stainless steel; CMF: cranio-maxillofacial).

| Device class | Mechanical requirements | Support period and resorption | Degradation target | Principal risks |
| --- | --- | --- | --- | --- |
| Screws, pins and K-wires | UTS >300 MPa, YS >200 MPa, elongation >15–20% [45,57,58]; pull-out 200–800 N (Mg) to ~3800 N (Fe–SS hybrid) [47,48]; insertion torque 0.5–3.5 Nm [48,70]; Mg pin bending strength 437–625 MPa [49,50] | Full support for 4–6 weeks; ≥50% strength at ~3 months; resorption within 9–24 months [41,51,53,54] | 0.2–0.5 mm yr⁻¹ [72] | Mg: too fast (>0.5–2.7 mm yr⁻¹); Fe: too slow (<0.1–0.16 mm yr⁻¹); pitting and corrosion fatigue in small sections [55,56,73] |
| Plates | Load-bearing: YS >200 MPa, UTS >300 MPa, elongation >10–18%, fatigue strength >100 MPa (10⁵–10⁷ cycles); non-load-bearing: YS ~100 MPa, UTS ~150 MPa, elongation >5% [40,41] | Full stability for 4–6 weeks; ≥50% strength at ~3 months; resorption within 9–24 months [39,40] | <0.5–1.0 mm yr⁻¹ [39,44] | Bending and loss of reduction, screw loosening, nonunion; hydrogen and alkalization (Mg) [44] |
| Intramedullary nails | YS >200 MPa, UTS >300 MPa, fatigue strength >100 MPa, elongation 10–20%, modulus 20–80 GPa [40,42,78]; 150,000–250,000 gait cycles (target ~500,000) and torque ≥2.3–5.5 Nm (ASTM F1264) [47,74,77] | Full capacity for 0–4 weeks; ≥50% strength at ~3 months; resorption within 12–24 months [40,53] | Not specified | Mg nails fail early (<70,000 cycles); marrow-fluid flow accelerates corrosion [74,81] |
| Porous scaffolds | Porosity ≥50% (70–90% favorable); pore size 300–600 μm; permeability 10⁻⁹–10⁻⁸ m²; interconnectivity 30–90% [76,84–90]; load-bearing: YS >200 MPa, UTS >300 MPa [40,41,45] | Degradation synchronized with bone ingrowth [85,91] | Not specified | Premature collapse (fast) or persistent stress shielding (slow); product retention [85,91] |
| CMF fixation | Non-load-bearing: modulus 10–40 GPa, YS >100 MPa, UTS >150 MPa, elongation >5%; mandibular: modulus 20–80 GPa, YS >200 MPa, UTS >300 MPa, fatigue >100 MPa, elongation >10% [40,92] | Full stability for 4–6 weeks; ≥30% (non-load-bearing) or ≥50% (mandibular) at ~3 months; resorption within 12–24 months [40,96,97] | Zn: 0.1–0.3 mm yr⁻¹ [40,57] | Mg: rapid corrosion and gas with limited soft-tissue cover; growth compatibility in pediatric patients [95,96] |
| Vascular stents | Radial strength 80–150 kPa (110–170 kPa for 3.0 × 18 mm coronary) [99,100]; recoil <15% (ISO 25539-2), target <4% [41]; fatigue durability 10–20 million cycles [41,57] | Full radial support for 3–6 months (up to 6–12 months); resorption within 12–24 months [29,41,102] | <~0.02 mm yr⁻¹ [41,57] | Pitting and stress-corrosion cracking at crowns; increased recoil with strut thinning [28,103] |

### Orthopedic Fixation Devices

Orthopedic fixation devices, including screws, pins, Kirschner wires, plates, and intramedullary nails, are intended to provide immediate mechanical stabilization after implantation, maintain structural integrity during the critical early stages of fracture repair, and progressively transfer load to the regenerating bone as healing advances [39,40]. Although their geometries and loading conditions differ considerably, these devices share a common biomechanical requirement: they must provide sufficient strength for temporary load bearing while minimizing mechanical mismatch with the surrounding skeletal tissue. Accordingly, biodegradable fixation systems are generally expected to achieve tensile and yield strengths exceeding approximately 300 MPa and 200 MPa, respectively, while maintaining elastic moduli closer to cortical bone (10-30 GPa) than conventional permanent implant materials [41,42]. This stiffness compatibility is particularly important because excessive implant rigidity can induce stress shielding, whereby the implant bears a disproportionate share of physiological loading, reducing mechanical stimulation of the adjacent bone and impairing remodeling and healing [40,43]. Among biodegradable metals, magnesium alloys exhibit the most favorable stiffness profile, with elastic moduli of approximately 41-45 GPa [44] whereas zinc alloys typically possess intermediate moduli of 90-110 GPa [45] and iron-based alloys exhibit substantially higher values of 190-211 GPa, making them more susceptible to stress-shielding effects despite their superior mechanical strength [46].

Within this shared framework, geometric design fundamentally influences the mechanical demands placed on each fixation device. Screws and interference screws are primarily evaluated through pull-out strength, insertion torque, and torsional resistance because their clinical performance depends on achieving sufficient anchorage within bone and resisting loosening under physiological loading before biological fixation develops. In contrast, pins and Kirschner wires have smaller cross-sectional dimensions and therefore lower structural reserves, making them particularly sensitive to the balance between strength and ductility. Plates and intramedullary nails operate under considerably higher absolute loading conditions and are consequently assessed using four-point bending and torsional fatigue testing rather than fixation strength alone.

Fixation stability in biodegradable orthopedic devices must evolve with fracture healing rather than be defined solely by initial mechanical performance. Across fixation systems, near-complete mechanical support is generally required during the first 4-6 weeks, corresponding to the inflammatory and soft-callus stages when the healing tissue contributes little intrinsic load-bearing capacity [53]. As hard-callus formation progresses between approximately 6 and 12 weeks, load-bearing devices are typically expected to retain at least 50% of their initial strength, whereas non-load-bearing applications may tolerate retention levels closer to 30% [41]. Following cortical union, which commonly occurs within 3-6 months, the implant assumes a progressively reduced mechanical role, with complete resorption ideally occurring within 9-24 months to enable full restoration of physiological load transfer [51,54]. Achieving this temporal balance is complicated by corrosion-induced degradation, which extends beyond uniform material loss. Localized pitting can generate stress concentrations that accelerate crack initiation and propagation, causing substantial reductions in ductility and fatigue resistance before equivalent losses in bulk strength are observed, while cyclic physiological loading can further accelerate corrosion relative to static conditions [55,56]. These challenges are particularly evident in magnesium-based fixation devices, where rapid degradation may compromise long-term mechanical retention; for example, LAE442 intramedullary pins lose approximately 40% of their bending strength within nine months in vivo and retain only about 12% of their original strength after 3.5 years, motivating extensive research into alloying and surface-modification strategies that can moderate degradation while preserving magnesium's favorable stiffness characteristics [51,54].

Orthopedic screws are among the most extensively investigated biodegradable metallic implants because they must provide immediate fixation stability, maintain mechanical support throughout bone healing, and gradually transfer load back to regenerating tissue as degradation progresses. Consequently, their design is governed by the coupled requirements of mechanical competence, fixation stability, degradation kinetics, and healing-time matching [40,45,57].

The mechanical performance of biodegradable screws is primarily determined by their strength, ductility, stiffness, and resistance to bending and shear loading. A commonly accepted benchmark for biodegradable metallic implants is an ultimate tensile strength (UTS) exceeding 300 MPa together with elongation at failure greater than 15–20% [45,57,58]. Advanced Zn-based alloys have achieved UTS values of 414–520 MPa with elongations of 26–84%, while Fe-based alloys can reach strengths of 400–900 MPa [40,58,59]. Furthermore, fixation screws must withstand bending and shear forces generated during physiological activities. Mg-Ca-Zn screws have demonstrated bending strengths ranging from 91 to 1384 MPa depending on core diameter and geometry, while metallic composite screw systems have tolerated flexural loads of approximately 260 N and double-shear loads of 2.2–2.5 kN [65,66].

Researchers commonly assess fixation performance through pull-out strength, insertion torque, torsional resistance, mechanical-retention capability, and degradation behavior. Pull-out testing performed according to ASTM F543 demonstrates substantial variation depending on anatomical application and bone quality [67]. ACL interference screws manufactured from biodegradable Mg alloys have reported pull-out strengths between 202.9 and 745.2 N, with fixation performance generally exceeding that of polymer-based alternatives [48,68]. For cortical and long-bone fixation, Mg screws have demonstrated pull-out forces of approximately 478 N, while experimental Fe-stainless steel hybrid screws have achieved values approaching 3800 N [47,69]. Insertion torque must be sufficiently high to ensure adequate thread engagement while remaining below stripping or fracture limits. Recommended insertion torques for orthopedic screws generally range from 0.5 to 3.5 Nm, with Mg interference screws typically ranging from 1.7 to 2.9 Nm [48,70]. As bone healing progresses, gradual strength loss becomes desirable because it promotes progressive load transfer from the implant to the recovering tissue [39,40]. To achieve this objective, an optimal degradation rate of approximately 0.2–0.5 mm year⁻¹ has been proposed [72]. Mg alloys commonly degrade faster than this target (>0.5–2.7 mm year⁻¹), so alloying and surface-modification strategies are needed to improve corrosion resistance [73,74]. Zn alloys exhibit more favorable degradation rates of 0.1–0.3 mm year⁻¹, closely matching the healing timeline, whereas pure Fe degrades too slowly (<0.1–0.16 mm year⁻¹) and may persist after functional healing is complete [46,73].

### Orthopedic Pins and Wires

Orthopedic pins and Kirschner wires (K-wires) are among the most demanding applications for biodegradable metallic implants because their small diameters provide limited mechanical reserve as degradation progresses. Unlike threaded screws, fixation stability comes primarily from the pin's bending and shear resistance, making these devices particularly sensitive to corrosion-induced cross-sectional loss. The design requirements are therefore governed by four interconnected factors: mechanical strength, stiffness compatibility, strength retention during degradation, and synchronization of implant lifetime with bone healing [42,64].

The mechanical requirements for biodegradable pins are largely dictated by physiological bending and shear loading. Magnesium-based pin systems generally satisfy these requirements, exhibiting bending strengths between approximately 360 and 625 MPa depending on alloy composition and processing route [49,65]. Iron-based systems can achieve substantially higher strengths, with hybrid Fe-stainless-steel designs exceeding 1.7 GPa in bending [47].

Because pins have limited cross-sectional area, maintaining mechanical integrity throughout degradation is often more critical than achieving exceptionally high initial strength. Corrosion-induced section loss, localized pitting, and corrosion-fatigue interactions progressively reduce bending stiffness and load-bearing capacity, particularly under cyclic physiological loading conditions [55,56]. Experimental studies of magnesium fixation pins have demonstrated substantial reductions in mechanical capacity during long-term implantation, with bending strength retention decreasing progressively as degradation proceeds [51,54]. These observations highlight the importance of controlling degradation kinetics through alloy design, surface modification, and coating technologies.

### Orthopedic Plates

Orthopedic bone plates are widely used for fracture fixation because they provide rigid stabilization, maintain anatomical alignment, and distribute physiological loads across the fracture site during healing. In biodegradable metallic systems, however, plate design is considerably more challenging than screw or pin fixation because plates are directly subjected to bending, torsional, and shear loads throughout healing while simultaneously degrading.

The mechanical requirements of biodegradable bone plates are generally more demanding than those of many other orthopedic fixation devices. For load-bearing applications such as long-bone fracture fixation, proposed design guidelines recommend a yield strength exceeding 200 MPa, an ultimate tensile strength greater than 300 MPa, elongation at failure above 10–18%, and fatigue strength exceeding 100 MPa over 105–107 loading cycles [40,41]. Beyond material properties, plate geometry strongly influences structural performance. Four-point bending analyses performed according to ASTM F382 show that magnesium alloys have substantially lower yield strength than conventional titanium alloys, causing bioresorbable magnesium locking compression plates to deform more under equivalent physiological loading conditions [39]. As a result, direct geometric replacement of conventional titanium plates is generally not feasible; instead, plate thickness, cross-sectional geometry, and structural topology must be optimized to maintain mechanical safety throughout degradation.

Unlike permanent fixation systems that permanently shield the fracture from mechanical loading, biodegradable plates create a time-dependent load-sharing relationship between the implant and the healing tissue [52]. Immediately after implantation, the fracture gap has negligible stiffness, so the plate carries nearly the entire physiological load. As mineralized callus forms and tissue stiffness increases during the repair phase, progressive implant degradation transfers load back to the regenerating bone [52]. Computational mechanoregulation studies have shown that this gradual transition can be advantageous because excessive fixation rigidity suppresses interfragmentary strain and reduces the mechanical stimuli required for endochondral ossification [39,52]. Conversely, controlled reductions in plate stiffness can promote cartilage formation, enhance callus maturation, and improve long-term remodeling outcomes. Therefore, biodegradable plate systems are increasingly viewed not as static fixation devices but as dynamic mechanical regulators that actively participate in healing.

A major motivation for biodegradable metallic bone plates is reducing stiffness mismatch between the implant and surrounding skeletal tissue. Conventional plate materials are substantially stiffer, with elastic moduli of approximately 189–213 GPa for 316L stainless steel, 220–240 GPa for Co-Cr alloys, and 110–127 GPa for Ti-6Al-4V [40]. This mismatch leads to stress shielding, whereby the implant carries a disproportionate fraction of the physiological load, reducing mechanical stimulation of the surrounding bone and promoting localized bone resorption, osteopenia, implant loosening, and increased refracture risk after hardware removal [42,44].

The degradation behavior of biodegradable plates directly influences fracture-healing outcomes because corrosion simultaneously alters implant mechanics and the local biological microenvironment. As corrosion progresses, reductions in plate thickness and cross-sectional area increase local stress concentrations and gradually reduce the load-bearing capability of the fixation construct [39,52]. Several studies suggest that degradation rates below approximately 0.5–1.0 mm yr⁻¹ are required to maintain sufficient structural support throughout healing, whereas excessively rapid corrosion may result in plate bending, loss of fracture reduction, screw loosening, or nonunion before adequate callus formation occurs [39,44]. In magnesium-based plates, rapid corrosion may further generate hydrogen gas accumulation and local alkalization, which can impair cell adhesion and compromise bone–implant integration unless suitable surface treatments or protective coatings are used [44].

### Intramedullary Nails

Intramedullary (IM) nails are among the most mechanically demanding biodegradable orthopedic implants because they are intended to fix long-bone fractures, including femoral and tibial diaphyseal fractures, where they experience continuous axial compression, bending, torsion, and cyclic fatigue loading during daily activities. Unlike plates and screws, which rely partly on cortical fixation, IM nails function as load-sharing devices positioned directly within the medullary canal and must maintain structural integrity throughout healing.

The mechanical performance of IM nails is commonly evaluated according to ASTM F1264, which specifies testing under cyclic four-point bending and torsional loading conditions representative of physiological gait. Under force-controlled fatigue testing, intramedullary fixation devices are expected to withstand approximately 150,000–250,000 gait cycles, corresponding to the first two to three months of fracture healing; incorporating a safety factor requires a target fatigue life approaching 500,000 cycles without failure [74,77]. Torsional stability is equally critical because rotational displacement can compromise fracture alignment and delay union. ASTM F1264 specifies minimum torsional torque capacities of approximately 2.3 Nm and 5.5 Nm for 3.5 mm and 5.0 mm diameter devices, respectively, with a minimum fracture angle of 180° [47].

The suitability of biodegradable metals for intramedullary fixation depends strongly on their balance between mechanical properties and degradation behavior. Magnesium alloys have elastic moduli of approximately 41–45 GPa, closely matching cortical bone and reducing stress shielding compared with titanium or stainless steel implants [62,74]. However, rapid corrosion often leads to premature fatigue degradation under load-bearing conditions. For example, LAE442 magnesium intramedullary nails subjected to cyclic bending in simulated physiological environments failed after fewer than 70,000 cycles, substantially below ASTM fatigue requirements [74]. Similarly, Wang et al. (2020) reported that more than 40% of uncoated JDBM magnesium nails implanted in rat femurs failed within four months, with implant breakage representing the dominant failure mode. These observations indicate that current uncoated magnesium alloys remain unsuitable as standalone fixation devices for heavily loaded long-bone fractures [44,53].

In contrast, zinc-based alloys have emerged as promising candidates for biodegradable intramedullary fixation because they offer a favorable compromise between strength and degradation rate. Advanced Zn alloys, including Zn-Mg-Mn and Zn-Li systems, have demonstrated yield strengths exceeding 200–300 MPa, ultimate tensile strengths of approximately 400–520 MPa, and elongations ranging from 26% to 84%, while maintaining relatively slow and uniform degradation rates of approximately 0.1–0.3 mm year⁻¹ [58,79]. These properties enable Zn alloys to overcome two major limitations of biodegradable metals: the rapid strength loss and hydrogen evolution associated with magnesium, and the excessively slow corrosion of iron. Iron-based alloys provide the highest mechanical strength among biodegradable metals, with ultimate tensile strengths of approximately 400–900 MPa and yield strengths reaching 650 MPa [40].

The degradation behavior of biodegradable IM nails is particularly important because fixation occurs within the medullary canal, where implants are continuously exposed to bone marrow and interstitial fluid flow. Physiological marrow fluid transport rates ranging from approximately 0.012 to 1.67 mL min⁻¹ can accelerate corrosion relative to static immersion conditions, intensifying the coupling between mechanical loading and degradation [81]. In magnesium-based systems, rapid degradation not only reduces structural integrity but also generates hydrogen gas and localized alkalization, potentially creating gas cavities that disrupt bone–implant contact, induce loosening, and compromise fracture stability [44,82].

To achieve these targets, considerable effort has focused on alloy optimization and surface engineering. Grain refinement through extrusion and thermomechanical processing can significantly enhance strength while improving corrosion uniformity [83]. Likewise, protective coatings such as fluoridation, plasma electrolytic oxidation, calcium-phosphate layers, and SrHPO₄ coatings have been shown to suppress early corrosion, reduce hydrogen evolution, and improve fatigue performance [44]. Hybrid fixation approaches, including magnesium devices used alongside stainless-steel tension bands or friction-welded Fe–SS systems, have also demonstrated improved load-bearing performance while retaining the benefits of biodegradability [47,53]. Collectively, these findings suggest that future biodegradable intramedullary nails will likely rely on corrosion-controlled alloys and multifunctional surface architectures that maintain mechanical competence throughout fracture healing while enabling complete long-term resorption.

### Porous Scaffolds

Porous biodegradable metallic scaffolds differ fundamentally from screws, pins, plates, and intramedullary nails because their primary function is not rigid fixation but the simultaneous provision of temporary mechanical support, tissue-guiding architecture, and a degradable microenvironment that promotes bone regeneration. Increasing porosity generally enhances cell infiltration and nutrient transport but inevitably reduces mechanical strength, making scaffold-architecture optimization a central challenge in developing biodegradable metallic implants [76,84,85].

Pore morphology and overall porosity strongly influence the biological performance of porous scaffolds. A minimum porosity of approximately 50% is generally required to facilitate meaningful bone ingrowth, whereas highly porous architectures with porosities between 70% and 90% provide more favorable conditions for cellular migration, extracellular matrix deposition, and tissue regeneration [84,85]. However, excessive porosity may substantially reduce load-bearing capacity because scaffold stiffness and strength decrease exponentially as relative density declines. Similarly, pore size plays a critical role in regulating cell behavior. Pore diameters below approximately 100 μm restrict osteoblast migration and vascular penetration, often resulting in fibrous tissue formation rather than direct bone regeneration [76,86]. In contrast, pore sizes between 300 and 600 μm are widely considered optimal for bone formation and vascular invasion because they provide sufficient space for mineralized tissue deposition while maintaining structural connectivity throughout the scaffold [84,87]. Furthermore, combining macropores that support tissue ingrowth with micropores that increase surface area enhances protein adsorption, ion exchange, and cell attachment, improving overall osteointegration [86].

Beyond pore size alone, successful bone regeneration requires adequate permeability and pore interconnectivity to support vascularization. Nutrient transport, oxygen delivery, and metabolic waste removal depend on physiological fluids circulating through the scaffold network. Additively manufactured metallic scaffolds have demonstrated permeability values comparable to trabecular bone, typically ranging from approximately 10⁻⁹ to 10⁻⁸ m² [88,89]. High permeability reduces the likelihood of hypoxic regions within large defects and promotes homogeneous tissue growth throughout the implant volume. Pore interconnectivity is equally important, with interconnected volume fractions of approximately 30–90% generally considered necessary for effective tissue invasion [90]. Interconnecting pore throats larger than 100 μm facilitate capillary infiltration and angiogenesis, whereas narrower channels may restrict vascular invasion and favor cartilaginous rather than osteogenic tissue development [76,90]. Consequently, modern computational scaffold design increasingly focuses on architected lattice structures such as gyroid, diamond, and triply periodic minimal surface (TPMS) geometries, which provide high permeability while preserving mechanical efficiency [88,89].

Unlike conventional fixation devices, scaffold performance is intrinsically linked to the evolving process of tissue regeneration. During the initial inflammatory and early osteogenic phases, the scaffold must maintain sufficient mechanical stability to prevent structural collapse. As mineralized bone gradually occupies the pore network, degradation-mediated reductions in stiffness become beneficial because they promote physiological load transfer and stimulate remodeling according to Wolff’s law [91]. An excessively rapid degradation rate may lead to premature loss of mechanical integrity and compromise defect bridging, whereas excessively slow degradation may result in persistent stress shielding and prolonged retention of degradation products [85,91]. Magnesium-based scaffolds release Mg²⁺ ions that promote osteogenesis and angiogenesis but can corrode rapidly and evolve hydrogen. Zinc-based scaffolds degrade more moderately and exhibit favorable biological responses without gas formation, while iron-based scaffolds provide superior mechanical strength but often require architectural design strategies to compensate for their slow degradation kinetics [89,91].

Mechanical compatibility remains an essential consideration despite the highly porous nature of scaffold structures. Dense metallic implants often produce large stiffness mismatches with native bone, whereas porous architecture substantially reduces the effective elastic modulus. Human cortical bone has an elastic modulus of approximately 10–30 GPa, while trabecular bone typically ranges from 0.1–5 GPa [40]. By controlling porosity and lattice topology, additively manufactured biodegradable metallic scaffolds can achieve mechanical properties within these ranges, reducing stress-shielding effects while maintaining adequate load-bearing capacity. Collectively, these requirements show that designing biodegradable metallic scaffolds is fundamentally an exercise in balancing architecture, mechanics, transport, and degradation to create a temporary three-dimensional environment that guides tissue regeneration before ultimately disappearing.

### Cranio-Maxillofacial (CMF) Fixation Devices

Cranio-maxillofacial (CMF) fixation devices encompass a diverse range of implants, including mini-plates, microplates, fixation screws, tacks, pins, and distractor systems used in cranial, orbital, midfacial, and mandibular reconstruction. Unlike orthopedic fixation devices designed for long bones, CMF implants operate in anatomically distinct environments, from minimally loaded cranial bones to the highly loaded mandibular skeleton. Consequently, mechanical requirements vary substantially by anatomical location: non-load-bearing cranial and midfacial applications prioritize dimensional stability and growth compatibility, whereas mandibular fixation requires greater strength and fatigue resistance to withstand cyclic masticatory loading [40,92].

For non-load-bearing applications such as cranial vault reconstruction, orbital fixation, and guided bone regeneration, biodegradable metallic implants are primarily required to maintain positional stability throughout early bone healing. Design guidelines generally recommend elastic moduli of 10–40 GPa, yield strengths above 100 MPa, ultimate tensile strengths above 150 MPa, and elongations above 5% [40]. These requirements provide adequate mechanical support while maintaining modulus compatibility with craniofacial bone, whose elastic modulus typically ranges from approximately 7–20 GPa in adults and varies substantially during pediatric growth [92]. In contrast, mandibular fixation devices must tolerate complex cyclic chewing forces and therefore require more demanding load-bearing criteria, typically including elastic moduli of 20–80 GPa, yield strengths above 200 MPa, ultimate tensile strengths exceeding 300 MPa, fatigue resistance greater than 100 MPa, and elongations above 10% [40,92].

At the device scale, fixation performance depends not only on bulk material properties but also on screw, tack, and pin stability. Recent magnesium membrane fixation systems have demonstrated substantially improved mechanical performance compared with conventional resorbable polymeric alternatives. For example, NOVAMag XS magnesium fixation screws exhibit higher pull-out resistance, shear strength, and torsional capacity than polymer-based tacks, enabling secure fixation of barrier membranes and bone graft materials while remaining fully resorbable [93]. Similarly, biomechanical studies of cranial distractor fixation have shown that contemporary bioresorbable pin systems provide sufficient pull-out and shear resistance for pediatric cranial vault distraction procedures, achieving fixation capacities substantially exceeding clinically required distraction forces [94].

The principal advantage of biodegradable CMF fixation is eliminating long-term hardware retention, which is particularly important in pediatric patients. Permanent titanium fixation can interfere with craniofacial growth, translocate within the developing skull, create palpable hardware, cause thermal sensitivity, and often require secondary removal procedures [95,96]. At approximately three months, non-load-bearing applications should retain at least 30% of their original strength, whereas mandibular fixation systems should retain at least 50% to prevent displacement during functional loading [40].

Among biodegradable metals, magnesium remains the most extensively investigated CMF material because of its favorable stiffness and osteogenic potential. However, rapid corrosion and hydrogen-gas evolution remain significant challenges, particularly in craniofacial regions where soft-tissue coverage is limited and cosmetic outcomes are critical (Byun et al., 2020; Li et al., 2024). Surface-treated and extrusion-processed Mg alloys, including WE43-based systems, have demonstrated improved corrosion control and fixation stability suitable for CMF applications (Byun et al., 2020). Zinc-based alloys offer an attractive alternative because they degrade more uniformly at approximately 0.1–0.3 mm year⁻¹ without generating gas cavities, while still providing adequate mechanical properties for low- to moderate-load applications [40,57]. Consequently, future CMF fixation systems will likely rely on carefully tailored degradation kinetics that preserve fixation during early healing, accommodate ongoing skeletal growth, and ultimately eliminate the need for secondary implant-removal procedures.

### Cardiovascular Stents

Biodegradable vascular stents represent one of the most demanding applications of biodegradable metals because they must provide sufficient radial support to maintain vessel patency while simultaneously undergoing controlled degradation in a dynamic hemodynamic environment. Unlike orthopedic implants, which primarily stabilize healing tissues, vascular stents serve as temporary scaffolds that prevent acute vessel recoil, support arterial remodeling, and disappear once vascular healing is complete.

The primary mechanical requirement of a biodegradable vascular stent is adequate radial strength to resist vessel collapse following balloon angioplasty. Current design benchmarks generally target radial compressive strengths of approximately 80–150 kPa, with 110–170 kPa considered suitable for conventional coronary stents near 3.0 mm × 18 mm [99,100]. Recent Zn-based stents have demonstrated particularly promising performance, with Zn-1.0Cu alloy designs achieving radial strengths around 120 kPa, comparable to commercial drug-eluting metallic stents and polymeric bioresorbable scaffolds [101]. Magnesium-based stents typically exhibit radial strengths between approximately 89 and 97 kPa, whereas nitrided iron-based scaffolds can achieve values exceeding 110 kPa while maintaining ultrathin strut geometries [99,100]. In addition to radial support, acute elastic recoil following balloon deflation must remain minimal to prevent restenosis and vessel renarrowing. Regulatory standards such as ISO 25539-2 generally require recoil values below 15%, although modern biodegradable vascular scaffolds increasingly target recoil values below 4% to approach the performance of permanent metallic stents [41].

Unlike permanent stents, biodegradable vascular scaffolds must retain mechanical function only during the period required for vascular repair. Clinical and experimental studies indicate that complete radial support is necessary during the first 3–6 months after implantation, when endothelial restoration, extracellular matrix remodeling, and vessel wall stabilization occur [29,102]. More complex lesions may require support for up to 6–12 months before complete load transfer to the healed artery becomes possible [102]. Following this period, the scaffold should gradually lose structural integrity and achieve near-complete bioabsorption within approximately 12–24 months, thereby restoring natural vasomotion and eliminating the long-term complications associated with permanent metallic implants [29,41]. To accomplish this objective, degradation rates must remain sufficiently low to prevent excessive ion release and premature loss of scaffolding capability, with corrosion penetration rates below approximately 0.02 mm year⁻¹ often cited as desirable targets for vascular applications [41,57].

Corrosion-induced degradation directly influences long-term scaffolding performance because reductions in strut thickness progressively decrease radial stiffness and increase the risk of structural instability. Finite-element studies have shown that uncontrolled degradation can increase radial recoil during service, leading to loss of lumen support before vascular remodeling is complete [28]. Localized pitting corrosion and stress-corrosion cracking are particularly problematic because stress concentrations frequently occur at crown junctions and hinge regions, where cyclic deformation is greatest. These regions can initiate crack growth and premature strut fracture [28,103]. Consequently, substantial effort has been devoted to structural optimization strategies, including rounded crown geometries, optimized strut topology, stress redistribution, and protective surface treatments, all of which promote more uniform degradation and prolong scaffold integrity.

Fatigue resistance represents an additional challenge because vascular stents experience continuous cyclic loading resulting from arterial pulsation and blood-pressure fluctuations. Permanent metallic stents are often designed to withstand hundreds of millions of loading cycles during long-term implantation; however, biodegradable stents need only survive throughout their intended support period. Current design guidelines generally recommend durability of approximately 10–20 million cycles without fracture, corresponding to the scaffold's expected mechanical-support lifetime [41,57]. Fatigue performance is commonly evaluated using pulsatile durability testing per ASTM F2477, while radial compression and recoil behavior are assessed per ASTM F3067 and ISO 25539-2 (ASTM F2477; ASTM F3067; ISO 25539-2).

## Time-dependent performance requirements

Writing prompt: Define application-specific trajectories for strength, stiffness, fixation stability, radial support, fatigue life, and acceptable degradation.

[Add review text and citations here]

## The healing-degradation performance corridor

Writing prompt: Introduce a quantitative safe window linking residual implant capacity, recovering tissue capacity, and required functional capacity.

[Add review text and citations here]

## Failure modes and acceptance criteria

Writing prompt: Cover premature section loss, localized corrosion, fatigue, loss of fixation, excessive persistence, and harmful accumulation of degradation products.

[Add review text and citations here]

## Standards as computational acceptance criteria

Writing prompt: Connect ASTM F3160, ASTM F3268, ISO 10993-15, device-specific mechanical tests, ASME V&V 40, and FDA computational credibility guidance to model inputs and outputs.

[Add review text and citations here]

# Design space of biodegradable metals

## Magnesium systems

Writing prompt: Summarize computationally relevant properties, rapid and localized corrosion, hydrogen evolution, film effects, and environmental sensitivity.

[Add review text and citations here]

## Zinc systems

Writing prompt: Discuss intermediate degradation behavior, oxygen-reduction control, corrosion-product films, localized attack, and the relative shortage of implant-scale mechanistic models.

[Add review text and citations here]

## Iron systems

Writing prompt: Discuss high strength, slow bioresorption, persistent products, and architecture or alloy strategies for accelerating controlled degradation.

[Add review text and citations here]

## Cross-material comparison

Writing prompt: Compare Mg, Zn, and Fe by degradation timescale, localization tendency, mechanical retention, dominant chemistry, and modelling maturity.

[Add review text and citations here]

## Recommended comparison table

Writing prompt: Insert a table comparing material advantage, dominant risk, mechanistic variables, calibration data, and computational design need.

[Add review text and citations here]

# Computational representations of biodegradation

## Empirical geometry-reduction models

Writing prompt: Review uniform surface recession, prescribed thinning, empirical mass-loss laws, and time-dependent geometry updates.

[Add review text and citations here]

## Continuum damage mechanics models

Writing prompt: Review scalar and nonlocal corrosion-damage variables, stiffness degradation, element deletion, mesh dependence, and VUMAT implementations.

[Add review text and citations here]

## Stochastic and cellular models

Writing prompt: Review Monte Carlo, cellular automata, probabilistic surface removal, pit nucleation, growth, and morphology statistics.

[Add review text and citations here]

## Mechanistic reaction-transport models

Writing prompt: Review anodic and cathodic kinetics, species transport, oxygen limitation, pH, buffering, film precipitation and dissolution, chloride effects, and flow.

[Add review text and citations here]

## Interface-tracking methods

Writing prompt: Compare element deletion, level-set, phase-field, ALE moving meshes, immersed methods, and voxel-based updates.

[Add review text and citations here]

## Model hierarchy and selection

Writing prompt: Explain when empirical, phenomenological, reduced mechanistic, or high-fidelity models are appropriate for screening, prediction, optimization, or regulatory evidence.

[Add review text and citations here]

# Coupling degradation with implant mechanics

## Sequential coupling

Writing prompt: Describe degradation prediction followed by geometry or material updating and structural analysis at selected time points.

[Add review text and citations here]

## Fully coupled mechanochemical degradation

Writing prompt: Discuss feedback among stress, strain, corrosion rate, section loss, stress redistribution, cracking, and fatigue.

[Add review text and citations here]

## Bone-implant and vessel-implant interaction

Writing prompt: Cover contact, friction, screw insertion and pretension, fracture-gap mechanics, callus maturation, balloon expansion, plaque contact, recoil, and vessel remodeling.

[Add review text and citations here]

## Residual structural performance metrics

Writing prompt: Compare residual strength, stiffness, fatigue resistance, pull-out force, interfragmentary motion, radial support, recoil, and lumen gain.

[Add review text and citations here]

## Coupling with tissue healing and remodelling

Writing prompt: Review models that jointly represent implant degradation, load transfer, tissue regeneration, bone ingrowth, and biological adaptation.

[Add review text and citations here]

# Geometry, architecture and manufacturing-aware design

## Conventional parameter optimization

Writing prompt: Review optimization of screw threads, core diameter, plate thickness, hole number, stent struts, connectors, and related geometric variables.

[Add review text and citations here]

## Porous and lattice structures

Writing prompt: Discuss porosity, pore size, unit-cell topology, anisotropy, surface-area-to-volume ratio, permeability, tissue ingrowth, and degradation localization.

[Add review text and citations here]

## Topology optimization

Writing prompt: Separate optimization at implantation, robust optimization across degraded states, and true time-dependent topology optimization.

[Add review text and citations here]

## Patient-specific design

Writing prompt: Cover CT segmentation, anatomical reconstruction, virtual implantation, physiological loading, printability, insertion-path checks, and surgeon verification.

[Add review text and citations here]

## Additive-manufacturing constraints

Writing prompt: Discuss minimum feature size, roughness, defects, residual stress, crystallographic texture, dimensional tolerance, and build orientation.

[Add review text and citations here]

## Coatings and spatially graded designs

Writing prompt: Review coating thickness, local surface modification, functionally graded porosity, and variable-stiffness architectures as design variables.

[Add review text and citations here]

# Optimization, surrogate models and uncertainty

## Design variables

Writing prompt: Organize material, kinetic, geometric, coating, architectural, environmental, and patient-specific variables.

[Add review text and citations here]

## Objectives

Writing prompt: Define objectives such as minimizing mass, stress shielding, peak stress, degradation non-uniformity, and harmful ion accumulation while preserving required support.

[Add review text and citations here]

## Constraints and multiobjective formulation

Writing prompt: State time-dependent mechanical, biological, manufacturing, and geometric constraints and explain Pareto trade-offs.

[Add review text and citations here]

## Sensitivity analysis and parameter identifiability

Writing prompt: Discuss local sensitivity, Morris elementary effects, Sobol indices, profile likelihood, parameter correlation, and practical identifiability.

[Add review text and citations here]

## Calibration and Bayesian inference

Writing prompt: Cover deterministic inverse fitting, Bayesian calibration, prior selection, posterior uncertainty, and transfer from coupon to implant geometries.

[Add review text and citations here]

## Surrogates and machine learning

Writing prompt: Review Gaussian processes, polynomial response surfaces, neural surrogates, reduced-order models, physics-informed methods, and active learning.

[Add review text and citations here]

## Robust and reliability-based design

Writing prompt: Discuss uncertainty propagation, probability of failure, worst-case design, reliability constraints, and patient-population variability.

[Add review text and citations here]

# Calibration, validation and model credibility

## Numerical verification

Writing prompt: Address code verification, solution verification, mesh and time-step convergence, conservation checks, and benchmark problems.

[Add review text and citations here]

## Experimental calibration data

Writing prompt: Review mass loss, hydrogen evolution, electrochemistry, surface recession, pH, ion concentration, film composition, and mechanical test data.

[Add review text and citations here]

## Local morphology validation

Writing prompt: Emphasize micro-CT, profilometry, pit-depth distributions, remaining-volume maps, and spatial comparison rather than mass loss alone.

[Add review text and citations here]

## Component-level and in vivo validation

Writing prompt: Discuss screw pull-out, plate bending, scaffold compression, stent radial testing, ex vivo morphology, animal data, and clinical observations.

[Add review text and citations here]

## Uncertainty quantification

Writing prompt: Separate parameter, experimental, model-form, geometric, loading, and biological uncertainties.

[Add review text and citations here]

## Context of use and credibility assessment

Writing prompt: Apply ASME V&V 40 and FDA risk-informed credibility concepts to define the question of interest, model risk, evidence requirements, and decision consequence.

[Add review text and citations here]

# Translational and regulatory considerations

## Role of computation in device development

Writing prompt: Explain how modelling supports design-space exploration, worst-case selection, test planning, interpolation, residual-strength prediction, and evidence integration.

[Add review text and citations here]

## Standards and regulatory pathways

Writing prompt: Discuss absorbable-metal material characterization, degradation testing, biological evaluation, device-specific mechanical standards, and jurisdiction-specific requirements.

[Add review text and citations here]

## Clinical and regulatory case studies

Writing prompt: Use approved or cleared biodegradable metallic devices, such as magnesium fixation screws and vascular scaffolds, to identify expected evidence packages.

[Add review text and citations here]

## Reproducibility and model reporting

Writing prompt: Recommend reporting governing equations, assumptions, parameters, calibration datasets, numerical settings, code availability, uncertainty, and validation limits.

[Add review text and citations here]

## Remaining barriers to translation

Writing prompt: Discuss in vitro-to-in vivo mismatch, limited longitudinal data, patient variability, manufacturing variability, computational cost, and absence of agreed digital endpoints.

[Add review text and citations here]

# Research gaps and future directions

## Lifetime-aware rather than initial-state optimization

Writing prompt: Explain the need to optimize the full support-degradation trajectory rather than only performance at implantation.

[Add review text and citations here]

## Transferable models for Zn and Fe

Writing prompt: Identify the need for material-specific but computationally comparable mechanistic frameworks beyond the Mg-dominated literature.

[Add review text and citations here]

## Coupled degradation-healing digital twins

Writing prompt: Discuss integration of corrosion, tissue regeneration, loading history, imaging, and patient-specific updating.

[Add review text and citations here]

## Multifidelity and data-efficient modelling

Writing prompt: Propose combining high-fidelity mechanistic models, reduced models, experiments, and in vivo data.

[Add review text and citations here]

## Standardized validation datasets and benchmarks

Writing prompt: Recommend open geometries, boundary conditions, degradation datasets, mechanical tests, and reporting conventions.

[Add review text and citations here]

## Regulatory-ready computational evidence

Writing prompt: Outline the path toward predefined contexts of use, credibility plans, traceable software, uncertainty bounds, and decision-relevant validation.

[Add review text and citations here]

# Conclusions

## Main synthesis

Writing prompt: Summarize how degradation physics, structural mechanics, biological response, geometry, and uncertainty must be treated as one design problem.

[Add review text and citations here]

## Design recommendations

Writing prompt: State the most defensible modelling and validation practices for future biodegradable metallic implant studies.

[Add review text and citations here]

## Final perspective

Writing prompt: Conclude with the role of computational design in moving biodegradable metals from material development toward predictable, patient-relevant, and clinically translatable devices.

[Add review text and citations here]

# Suggested figures

[Add and update content here]

# Suggested tables

[Add and update content here]

# References

[1]	M.S. Baltatu, C.A. Tugui, M.C. Perju, M. Benchea, M.C. Spataru, A.V. Sandu, P. Vizureanu, Biocompatible Titanium Alloys used in Medical Applications, Rev. Chim. 70 (2019) 1302–1306. https://doi.org/10.37358/RC.19.4.7114.

[2]	K.-H. Frosch, K.M. Stürmer, Metallic Biomaterials in Skeletal Repair, Eur. J. Trauma 32 (2006) 149–159. https://doi.org/10.1007/s00068-006-6041-1.

[3]	W. Abd-Elaziem, M.A. Darwish, A. Hamada, W.M. Daoush, Titanium-Based alloys and composites for orthopedic implants Applications: A comprehensive review, Mater. Des. 241 (2024) 112850. https://doi.org/10.1016/j.matdes.2024.112850.

[4]	G. Mani, D. Porter, S. Collins, T. Schatz, A. Ornberg, R. Shulfer, A review on manufacturing processes of cobalt‐chromium alloy implants and its impact on corrosion resistance and biocompatibility, J. Biomed. Mater. Res. B Appl. Biomater. 112 (2024) e35431. https://doi.org/10.1002/jbm.b.35431.

[5]	M. Niinomi, M. Nakai, J. Hieda, Development of new metallic alloys for biomedical applications, Acta Biomater. 8 (2012) 3888–3903. https://doi.org/10.1016/j.actbio.2012.06.037.

[6]	M. Kaveh, M. Badrossamay, E. Foroozmehr, M. Kharaziha, Biodegradability, biocompatibility, and mechanical behavior of additively manufactured zinc scaffolds, J. Mech. Behav. Biomed. Mater. 163 (2025) 106868. https://doi.org/10.1016/j.jmbbm.2024.106868.

[7]	S. Limmahakhun, A. Oloyede, K. Sitthiseripratip, Y. Xiao, C. Yan, Stiffness and strength tailoring of cobalt chromium graded cellular structures for stress-shielding reduction, Mater. Des. 114 (2017) 633–641. https://doi.org/10.1016/j.matdes.2016.11.090.

[8]	X.P. Tan, Y.J. Tan, C.S.L. Chow, S.B. Tor, W.Y. Yeong, Metallic powder-bed based 3D printing of cellular scaffolds for orthopaedic implants: A state-of-the-art review on manufacturing, topological design, mechanical properties and biocompatibility, Mater. Sci. Eng. C 76 (2017) 1328–1343. https://doi.org/10.1016/j.msec.2017.02.094.

[9]	P. Erne, M. Schier, T.J. Resink, The Road to Bioabsorbable Stents: Reaching Clinical Reality?, Cardiovasc. Intervent. Radiol. 29 (2006) 11–16. https://doi.org/10.1007/s00270-004-0341-9.

[10]	R. Erbel, C. Di Mario, J. Bartunek, J. Bonnier, B. de Bruyne, F.R. Eberli, P. Erne, M. Haude, B. Heublein, M. Horrigan, C. Ilsley, D. Böse, J. Koolen, T.F. Lüscher, N. Weissman, R. Waksman, Temporary scaffolding of coronary arteries with bioabsorbable magnesium stents: a prospective, non-randomised multicentre trial, The Lancet 369 (2007) 1869–1875. https://doi.org/10.1016/S0140-6736(07)60853-8.

[11]	M. Hussain, S.M. Khan, M. Shafiq, N. Abbas, U. Sajjad, K. Hamid, Advances in biodegradable materials: Degradation mechanisms, mechanical properties, and biocompatibility for orthopedic applications, Heliyon 10 (2024) e32713. https://doi.org/10.1016/j.heliyon.2024.e32713.

[12]	N.S. Manam, W.S.W. Harun, D.N.A. Shri, S.A.C. Ghani, T. Kurniawan, M.H. Ismail, M.H.I. Ibrahim, Study of corrosion in biocompatible metals for implants: A review, J. Alloys Compd. 701 (2017) 698–715. https://doi.org/10.1016/j.jallcom.2017.01.196.

[13]	J. Venezuela, M.S. Dargusch, The influence of alloying and fabrication techniques on the mechanical properties, biodegradability and biocompatibility of zinc: A comprehensive review, Acta Biomater. 87 (2019) 1–40. https://doi.org/10.1016/j.actbio.2019.01.035.

[14]	L. Tan, X. Yu, P. Wan, K. Yang, Biodegradable Materials for Bone Repairs: A Review, J. Mater. Sci. Technol. 29 (2013) 503–513. https://doi.org/10.1016/j.jmst.2013.03.002.

[15]	S. Zhao, M. Tayyebi, Mahdireza Yarigarravesh, G. Hu, A review of magnesium corrosion in bio-applications: mechanism, classification, modeling, in-vitro, and in-vivo experimental testing, and tailoring Mg corrosion rate, J. Mater. Sci. 58 (2023) 12158–12181. https://doi.org/10.1007/s10853-023-08782-z.

[16]	Y. Li, H. Jahr, K. Lietaert, P. Pavanram, A. Yilmaz, L.I. Fockaert, M.A. Leeflang, B. Pouran, Y. Gonzalez-Garcia, H. Weinans, J.M.C. Mol, J. Zhou, A.A. Zadpoor, Additively manufactured biodegradable porous iron, Acta Biomater. 77 (2018) 380–393. https://doi.org/10.1016/j.actbio.2018.07.011.

[17]	J. Rao, H. Gao, J. Sun, R. Yu, D. Zhao, Y. Ding, A Critical Review of Biodegradable Zinc Alloys toward Clinical Applications, ACS Biomater. Sci. Eng. 10 (2024) 5454–5473. https://doi.org/10.1021/acsbiomaterials.4c00210.

[18]	G. Katarivas Levy, J. Goldman, E. Aghion, The Prospects of Zinc as a Structural Material for Biodegradable Implants—A Review Paper, Metals 7 (2017) 402. https://doi.org/10.3390/met7100402.

[19]	H. Kabir, K. Munir, C. Wen, Y. Li, Recent research and progress of biodegradable zinc alloys and composites for biomedical applications: Biomechanical and biocorrosion perspectives, Bioact. Mater. 6 (2021) 836–879. https://doi.org/10.1016/j.bioactmat.2020.09.013.

[20]	Y. Li, J. Zhou, P. Pavanram, M.A. Leeflang, L.I. Fockaert, B. Pouran, N. Tümer, K.-U. Schröder, J.M.C. Mol, H. Weinans, H. Jahr, A.A. Zadpoor, Additively manufactured biodegradable porous magnesium, Acta Biomater. 67 (2018) 378–392. https://doi.org/10.1016/j.actbio.2017.12.008.

[21]	N.T. Kirkland, N. Birbilis, M.P. Staiger, Assessing the corrosion of biodegradable magnesium implants: A critical review of current methodologies and their limitations, Acta Biomater. 8 (2012) 925–936. https://doi.org/10.1016/j.actbio.2011.11.014.

[22]	Y.F. Zheng, X.N. Gu, F. Witte, Biodegradable metals, Mater. Sci. Eng. R Rep. 77 (2014) 1–34. https://doi.org/10.1016/j.mser.2014.01.001.

[23]	M. Prakasam, J. Locs, K. Salma-Ancane, D. Loca, A. Largeteau, L. Berzina-Cimdina, Biodegradable Materials and Metallic Implants—A Review, J. Funct. Biomater. 8 (2017) 44. https://doi.org/10.3390/jfb8040044.

[24]	A. Joshi, G. Dias, M.P. Staiger, In silico modelling of the corrosion of biodegradable magnesium-based biomaterials: modelling approaches, validation and future perspectives, Biomater. Transl. 2 (2021) 257–271. https://doi.org/10.12336/biomatertransl.2021.03.008.

[25]	J.A. Grogan, B.J. O’Brien, S.B. Leen, P.E. McHugh, A corrosion model for bioabsorbable metallic stents, Acta Biomater. 7 (2011) 3523–3533. https://doi.org/10.1016/j.actbio.2011.05.032.

[26]	D. Gastaldi, V. Sassi, L. Petrini, M. Vedani, S. Trasatti, F. Migliavacca, Continuum damage model for bioresorbable magnesium alloy devices — Application to coronary stents, J. Mech. Behav. Biomed. Mater. 4 (2011) 352–365. https://doi.org/10.1016/j.jmbbm.2010.11.003.

[27]	F. Saconi, G.H. Diaz, A.C. Vieira, M.L. Ribeiro, Experimental Characterization and Numerical Modeling of the Corrosion Effect on the Mechanical Properties of the Biodegradable Magnesium Alloy WE43 for Orthopedic Applications, Materials 15 (2022) 7164. https://doi.org/10.3390/ma15207164.

[28]	K. Peng, X. Cui, A. Qiao, Y. Mu, Mechanical analysis of a novel biodegradable zinc alloy stent based on a degradation model, Biomed. Eng. OnLine 18 (2019) 39. https://doi.org/10.1186/s12938-019-0661-2.

[29]	Y. Li, H. Jahr, J. Zhou, A.A. Zadpoor, Additively manufactured biodegradable porous metals, Acta Biomater. 115 (2020) 29–50. https://doi.org/10.1016/j.actbio.2020.08.018.

[30]	E.B. Montufar, Bone Response to Biodegradable Metals and In Vitro Evaluation of the Cytocompatibility, JOM 77 (2025) 4473–4492. https://doi.org/10.1007/s11837-025-07353-8.

[31]	J. Zhang, Y. Jiang, Z. Shang, B. Zhao, M. Jiao, W. Liu, M. Cheng, B. Zhai, Y. Guo, B. Liu, X. Shi, B. Ma, Biodegradable metals for bone defect repair: A systematic review and meta-analysis based on animal studies, Bioact. Mater. 6 (2021) 4027–4052. https://doi.org/10.1016/j.bioactmat.2021.03.035.

[32]	Z. Wang, B. Liu, B. Yin, Y. Zheng, Y. Tian, P. Wen, Comprehensive review of additively manufactured biodegradable magnesium implants for repairing bone defects from biomechanical and biodegradable perspectives, Front. Chem. 10 (2022) 1066103. https://doi.org/10.3389/fchem.2022.1066103.

[33]	N. Pohl, D. Priebe, T. AlBaraghtheh, S. Schimek, D.C.F. Wieland, D. Krüger, S. Trostorff, R. Willumeit-Römer, R. Köhl, B. Zeller-Plumhoff, Computational modelling of bone growth and mineralization surrounding biodegradable Mg-based and permanent Ti implants, Comput. Struct. Biotechnol. J. 27 (2025) 1394–1406. https://doi.org/10.1016/j.csbj.2025.02.027.

[34]	H. Mndlovu, P. Kumar, L.C. Du Toit, Y.E. Choonara, A review of biomaterial degradation assessment approaches employed in the biomedical field, Npj Mater. Degrad. 8 (2024) 66. https://doi.org/10.1038/s41529-024-00487-1.

[35]	C. Panaghie, G. Zegan, A. Sodor, N. Cimpoeșu, N.-M. Lohan, B. Istrate, A.-M. Roman, N. Ioanid, Analysis of Degradation Products of Biodegradable ZnMgY Alloy, Materials 16 (2023) 3092. https://doi.org/10.3390/ma16083092.

[36]	T. AlBaraghtheh, R. Willumeit‐Römer, B. Zeller‐Plumhoff, Best Practices in Developing a Workflow for Uncertainty Quantification for Modeling the Biodegradation of Mg‐Based Implants, Adv. Sci. 11 (2024) 2403543. https://doi.org/10.1002/advs.202403543.

[37]	A. Amerinatanzi, R. Mehrabi, H. Ibrahim, A. Dehghan, N. Shayesteh Moghaddam, M. Elahinia, Predicting the Biodegradation of Magnesium Alloy Implants: Modeling, Parameter Identification, and Validation, Bioengineering 5 (2018) 105. https://doi.org/10.3390/bioengineering5040105.

[38]	M. Barzegari, D. Mei, S.V. Lamaka, L. Geris, Computational modeling of degradation process of biodegradable magnesium biomaterials, Corros. Sci. 190 (2021) 109674. https://doi.org/10.1016/j.corsci.2021.109674.

[39]	G. Chandra, A. Pandey, Design approaches and challenges for biodegradable bone implants: a review, Expert Rev. Med. Devices 18 (2021) 629–647. https://doi.org/10.1080/17434440.2021.1935875.

[40]	A. Cherqaoui, C. Paternoster, D. Mantovani, Biodegradable Metals and Corrosion Control: Challenges, Limits and New Opportunities for Innovating in Orthopedic Fixations, Materials 19 (2026) 1789. https://doi.org/10.3390/ma19091789.

[41]	P.K. Bowen, E.R. Shearier, S. Zhao, R.J. Guillory, F. Zhao, J. Goldman, J.W. Drelich, Biodegradable Metals for Cardiovascular Stents: from Clinical Concerns to Recent Zn‐Alloys, Adv. Healthc. Mater. 5 (2016) 1121–1140. https://doi.org/10.1002/adhm.201501019.

[42]	Z. Sheikh, S. Najeeb, Z. Khurshid, V. Verma, H. Rashid, M. Glogauer, Biodegradable Materials for Bone Repair and Tissue Engineering Applications, Materials 8 (2015) 5744–5794. https://doi.org/10.3390/ma8095273.

[43]	M. Asgari, R. Hang, C. Wang, Z. Yu, Z. Li, Y. Xiao, Biodegradable Metallic Wires in Dental and Orthopedic Applications: A Review, Metals 8 (2018) 212. https://doi.org/10.3390/met8040212.

[44]	Z. Wang, X. Wang, J. Pei, Y. Tian, J. Zhang, C. Jiang, J. Huang, Z. Pang, Y. Cao, X. Wang, S. An, X. Wang, H. Huang, G. Yuan, Z. Yan, Degradation and osteogenic induction of a SrHPO4-coated Mg–Nd–Zn–Zr alloy intramedullary nail in a rat femoral shaft fracture model, Biomaterials 247 (2020) 119962. https://doi.org/10.1016/j.biomaterials.2020.119962.

[45]	Y. Liu, T. Du, A. Qiao, Y. Mu, H. Yang, Zinc-Based Biodegradable Materials for Orthopaedic Internal Fixation, J. Funct. Biomater. 13 (2022) 164. https://doi.org/10.3390/jfb13040164.

[46]	R. Gorejová, L. Haverová, R. Oriňaková, A. Oriňak, M. Oriňak, Recent advancements in Fe-based biodegradable materials for bone repair, J. Mater. Sci. 54 (2019) 1913–1947. https://doi.org/10.1007/s10853-018-3011-z.

[47]	A.K. Nasution, M.F. Ulum, M.R.A. Kadir, H. Hermawan, Mechanical and corrosion properties of partially degradable bone screws made of pure iron and stainless steel 316L by friction welding, Sci. China Mater. 61 (2018) 593–606. https://doi.org/10.1007/s40843-017-9057-3.

[48]	N. Siroros, R. Merfort, Y. Liu, M. Praster, F. Migliorini, N. Maffulli, R. Michalik, F. Hildebrand, J. Eschweiler, Mechanical properties of a bioabsorbable magnesium interference screw for anterior cruciate ligament reconstruction in various testing bone materials, Sci. Rep. 13 (2023) 12342. https://doi.org/10.1038/s41598-023-39513-8.

[49]	M. Thomann, Ch. Krause, D. Bormann, N. von der Höh, H. Windhagen, A. Meyer‐Lindenberg, Comparison of the resorbable magnesium . alloys LAE442 und MgCa0.8 concerning their mechanical properties, their progress of degradation and the bone‐implant‐contact after 12 months implantation duration in a rabbit model, Mater. Werkst. 40 (2009) 82–87. https://doi.org/10.1002/mawe.200800412.

[50]	T. POHJONEN, P. HELEVIRTA, P. TORMALA, K. KOSKIKARE, H. PATIALA, P. ROKKANEN, Strength retention of self-reinforced poly-L-lactide screws. A comparison of compression moulded and machine cut screws, J. Mater. Sci. Mater. Med. 8 (1997) 311–320. https://doi.org/10.1023/A:1018516529591.

[51]	H. Waizy, J.-M. Seitz, J. Reifenrath, A. Weizbauer, F.-W. Bach, A. Meyer-Lindenberg, B. Denkena, H. Windhagen, Biodegradable magnesium implants for orthopedic applications, J. Mater. Sci. 48 (2013) 39–50. https://doi.org/10.1007/s10853-012-6572-2.

[52]	A. Vautrin, M. Wesseling, R. Wirix-Speetjens, M.J. Gomez-Benito, Time-dependent in silico modelling of orthognathic surgery to support the design of biodegradable bone plates, J. Mech. Behav. Biomed. Mater. 121 (2021) 104641. https://doi.org/10.1016/j.jmbbm.2021.104641.

[53]	D.H.K. Chow, J. Wang, P. Wan, L. Zheng, M.T.Y. Ong, L. Huang, W. Tong, L. Tan, K. Yang, L. Qin, Biodegradable magnesium pins enhanced the healing of transverse patellar fracture in rabbits, Bioact. Mater. 6 (2021) 4176–4185. https://doi.org/10.1016/j.bioactmat.2021.03.044.

[54]	N. Angrisani, J. Reifenrath, F. Zimmermann, R. Eifler, A. Meyer-Lindenberg, K. Vano-Herrera, C. Vogt, Biocompatibility and degradation of LAE442-based magnesium alloys after implantation of up to 3.5 years in a rabbit model, Acta Biomater. 44 (2016) 355–365. https://doi.org/10.1016/j.actbio.2016.08.002.

[55]	Y. Koo, H.-B. Lee, Z. Dong, R. Kotoka, J. Sankar, N. Huang, Y. Yun, The Effects of Static and Dynamic Loading on Biodegradable Magnesium Pins In Vitro and In Vivo, Sci. Rep. 7 (2017). https://doi.org/10.1038/s41598-017-14836-5.

[56]	S. Amukarimi, M. Mozafari, Biodegradable magnesium‐based biomaterials: An overview of challenges and opportunities, MedComm 2 (2021) 123–144. https://doi.org/10.1002/mco2.59.

[57]	Y. Liu, B. Lu, Z. Cai, Recent Progress on Mg- and Zn-Based Alloys for Biodegradable Vascular Stent Applications, J. Nanomater. 2019 (2019) 1–16. https://doi.org/10.1155/2019/1310792.

[58]	C. Lu, C. Song, Y. Yu, L. Yang, W. Zheng, F. Luo, Y. Xiao, J. Luo, J. Xu, Biodegradable zinc alloys with high strength and suitable mechanical integrity as bone repair metals, Sci. Rep. 14 (2024) 30558. https://doi.org/10.1038/s41598-024-78842-0.

[59]	H. Yang, B. Jia, Z. Zhang, X. Qu, G. Li, W. Lin, D. Zhu, K. Dai, Y. Zheng, Alloying design of biodegradable zinc as promising bone implants for load-bearing applications, Nat. Commun. 11 (2020) 401. https://doi.org/10.1038/s41467-019-14153-7.

[60]	M. Yazdimamaghani, M. Razavi, D. Vashaee, K. Moharamzadeh, A.R. Boccaccini, L. Tayebi, Porous magnesium-based scaffolds for tissue engineering, Mater. Sci. Eng. C 71 (2017) 1253–1266. https://doi.org/10.1016/j.msec.2016.11.027.

[61]	C. Shuai, S. Li, S. Peng, P. Feng, Y. Lai, C. Gao, Biodegradable metallic bone implants, Mater. Chem. Front. 3 (2019) 544–562. https://doi.org/10.1039/C8QM00507A.

[62]	I. Antoniac, M. Miculescu, V. Mănescu (Păltânea), A. Stere, P.H. Quan, G. Păltânea, A. Robu, K. Earar, Magnesium-Based Alloys Used in Orthopedic Surgery, Materials 15 (2022) 1148. https://doi.org/10.3390/ma15031148.

[63]	A. Kansal, N. Dixit, S. Kumar, A. Dvivedi, P. Kuamr, Review: zinc alloys for biomedical applications—alloy design, processing techniques, and bio-functional performance, J. Mater. Sci. 61 (2026) 3676–3713. https://doi.org/10.1007/s10853-025-12110-y.

[64]	H. Yang, B. Jia, Z. Zhang, X. Qu, G. Li, W. Lin, D. Zhu, K. Dai, Y. Zheng, Alloying design of biodegradable zinc as promising bone implants for load-bearing applications, Nat. Commun. 11 (2020) 401. https://doi.org/10.1038/s41467-019-14153-7.

[65]	S.Y. Cho, S. Chae, K.W. Choi, H.K. Seok, Y.C. Kim, J.Y. Jung, S.J. Yang, G.J. Kwon, J.T. Kim, M. Assad, Biocompatibility and strength retention of biodegradable Mg‐Ca‐Zn alloy bone implants, J. Biomed. Mater. Res. B Appl. Biomater. 101B (2013) 201–212. https://doi.org/10.1002/jbm.b.32813.

[66]	R.M. Felfel, I. Ahmed, A.J. Parsons, C.D. Rudd, Bioresorbable composite screws manufactured via forging process: Pull-out, shear, flexural and degradation characteristics, J. Mech. Behav. Biomed. Mater. 18 (2013) 108–122. https://doi.org/10.1016/j.jmbbm.2012.11.009.

[67]	F04 Committee, Specification and Test Methods for Absorbable Plates and Screws for Internal Fixation Implants, (n.d.). https://doi.org/10.1520/F2502-24.

[68]	N. Siroros, R. Merfort, Y. Liu, M. Praster, F. Hildebrand, R. Michalik, J. Eschweiler, In Vitro Investigation of the Fixation Performance of a Bioabsorbable Magnesium ACL Interference Screw Compared to a Conventional Interference Screw, Life 13 (2023) 484. https://doi.org/10.3390/life13020484.

[69]	J.-W. Li, C.-F. Du, C.-X. Yuchi, C.-Q. Zhang, Application of Biodegradable Materials in Orthopedics, J. Med. Biol. Eng. 39 (2019) 633–645. https://doi.org/10.1007/s40846-019-00469-8.

[70]	N. Siroros, R. Merfort, Y. Liu, M. Praster, F. Hildebrand, R. Michalik, J. Eschweiler, In Vitro Investigation of the Fixation Performance of a Bioabsorbable Magnesium ACL Interference Screw Compared to a Conventional Interference Screw, Life 13 (2023) 484. https://doi.org/10.3390/life13020484.

[71]	J. Henkel, M.A. Woodruff, D.R. Epari, R. Steck, V. Glatt, I.C. Dickinson, P.F.M. Choong, M.A. Schuetz, D.W. Hutmacher, Bone Regeneration Based on Tissue Engineering Conceptions — A 21st Century Perspective, Bone Res. 1 (2013) 216–248. https://doi.org/10.4248/BR201303002.

[72]	C. Shuai, S. Li, S. Peng, P. Feng, Y. Lai, C. Gao, Biodegradable metallic bone implants, Mater. Chem. Front. 3 (2019) 544–562. https://doi.org/10.1039/C8QM00507A.

[73]	M. Pogorielov, E. Husak, A. Solodivnik, S. Zhdanov, Magnesium-based biodegradable alloys: Degradation, application, and alloying elements, Interv. Med. Appl. Sci. 9 (2017) 27–38. https://doi.org/10.1556/1646.9.2017.1.04.

[74]	M. Krämer, M. Schilling, R. Eifler, B. Hering, J. Reifenrath, S. Besdo, H. Windhagen, E. Willbold, A. Weizbauer, Corrosion behavior, biocompatibility and biomechanical stability of a prototype magnesium-based biodegradable intramedullary nailing system, Mater. Sci. Eng. C 59 (2016) 129–135. https://doi.org/10.1016/j.msec.2015.10.006.

[75]	H. Wang, Z. Hao, S. Wen, Do biodegradable magnesium alloy intramedullary interlocking nails prematurely lose fixation stability in the treatment of tibial fracture? A numerical simulation, J. Mech. Behav. Biomed. Mater. 65 (2017) 117–126. https://doi.org/10.1016/j.jmbbm.2016.08.014.

[76]	V. Karageorgiou, D. Kaplan, Porosity of 3D biomaterial scaffolds and osteogenesis, Biomaterials 26 (2005) 5474–5491. https://doi.org/10.1016/j.biomaterials.2005.02.002.

[77]	F04 Committee, Specification and Test Methods for Intramedullary Fixation Devices, (n.d.). https://doi.org/10.1520/F1264-24.

[78]	N. Yang, N. Balasubramani, J. Venezuela, H. Bielefeldt-Ohmann, R. Allavena, S. Almathami, M. Dargusch, Microstructure refinement in biodegradable Zn-Cu-Ca alloy for enhanced mechanical properties, degradation homogeneity, and strength retention in simulated physiological condition, J. Mater. Sci. Technol. 125 (2022) 1–14. https://doi.org/10.1016/j.jmst.2022.03.006.

[79]	D. Shen, Y. Li, J. Shi, T. Zhang, J.-J. Nie, D. Chen, D. Xia, Y. Zheng, Biodegradable Zn-Li-Mn alloy to achieve optimal strength and ductility for bone implants, Acta Biomater. 199 (2025) 483–499. https://doi.org/10.1016/j.actbio.2025.04.056.

[80]	M. Peuster, P. Wohlsein, M. Brügmann, M. Ehlerding, K. Seidler, C. Fink, H. Brauer, A. Fischer, G. Hausdorf, A novel approach to temporary stenting: degradable cardiovascular stents produced from corrodible metal—results 6–18 months after implantation into New Zealand white rabbits, Heart 86 (2001) 563–569. https://doi.org/10.1136/heart.86.5.563.

[81]	A.P. Md. Saad, N. Jasmawati, M.N. Harun, M.R. Abdul Kadir, H. Nur, H. Hermawan, A. Syahrom, Dynamic degradation of porous magnesium under a simulated environment of human cancellous bone, Corros. Sci. 112 (2016) 495–506. https://doi.org/10.1016/j.corsci.2016.08.017.

[82]	J. Ye, S. Bai, X. He, J. Chen, C. Xie, Effects of Orientations, Roughnesses, and Cavities on Stress-Corrosion Coupled Damage in Magnesium, Crystals 12 (2022) 635. https://doi.org/10.3390/cryst12050635.

[83]	C. Rössig, N. Angrisani, P. Helmecke, S. Besdo, J.-M. Seitz, B. Welke, N. Fedchenko, H. Kock, J. Reifenrath, In vivo evaluation of a magnesium-based degradable intramedullary nailing system in a sheep model, Acta Biomater. 25 (2015) 369–383. https://doi.org/10.1016/j.actbio.2015.07.025.

[84]	Y. Jin, J. Li, H. Fan, J. Du, Y. He, Biomechanics and Mechanobiology of Additively Manufactured Porous Load‐Bearing Bone Implants, Small 21 (2025) 2409955. https://doi.org/10.1002/smll.202409955.

[85]	A. Motaharinia, J.W. Drelich, S. Sharif, A.F. Ismail, F. Naeimi, A. Glover, M. Ebrahiminejad, H.R. Bakhsheshi-Rad, Overview of porous magnesium-based scaffolds: development, properties and biomedical applications, Mater. Futur. 4 (2025) 012401. https://doi.org/10.1088/2752-5724/ad9493.

[86]	N. Abbasi, S. Hamlet, R.M. Love, N.-T. Nguyen, Porous scaffolds for bone regeneration, J. Sci. Adv. Mater. Devices 5 (2020) 1–9. https://doi.org/10.1016/j.jsamd.2020.01.007.

[87]	M. Kumar, S.S. Mohol, V. Sharma, A computational approach from design to degradation of additively manufactured scaffold for bone tissue engineering application, Rapid Prototyp. J. 28 (2022) 1956–1967. https://doi.org/10.1108/RPJ-12-2021-0336.

[88]	Y. Li, H. Jahr, P. Pavanram, F.S.L. Bobbert, U. Puggi, X.-Y. Zhang, B. Pouran, M.A. Leeflang, H. Weinans, J. Zhou, A.A. Zadpoor, Additively manufactured functionally graded biodegradable porous iron, Acta Biomater. 96 (2019) 646–661. https://doi.org/10.1016/j.actbio.2019.07.013.

[89]	W. Song, D. Zhao, F. Guo, J. Wang, Y. Wang, X. Wang, Z. Han, W. Fan, Y. Liu, Z. Xu, L. Chen, Additive manufacturing of degradable metallic scaffolds for material-structure-driven diabetic maxillofacial bone regeneration, Bioact. Mater. 36 (2024) 413–426. https://doi.org/10.1016/j.bioactmat.2024.06.028.

[90]	H. Mohammadi, M. Sepantafar, N. Muhamad, A. Bakar Sulong, How Does Scaffold Porosity Conduct Bone Tissue Regeneration?, Adv. Eng. Mater. 23 (2021) 2100463. https://doi.org/10.1002/adem.202100463.

[91]	M. Salama, M.F. Vaz, R. Colaço, C. Santos, M. Carmezim, Biodegradable Iron and Porous Iron: Mechanical Properties, Degradation Behaviour, Manufacturing Routes and Biomedical Applications, J. Funct. Biomater. 13 (2022) 72. https://doi.org/10.3390/jfb13020072.

[92]	P. Li, J. Dai, Y. Li, D. Alexander, J. Čapek, J. Geis-Gerstorfer, G. Wan, J. Han, Z. Yu, A. Li, Zinc based biodegradable metals for bone repair and regeneration: Bioactivity and molecular mechanisms, Mater. Today Bio 25 (2024) 100932. https://doi.org/10.1016/j.mtbio.2023.100932.

[93]	T. Franke, T. Korzinskas, Guided Bone Regeneration in the Posterior Mandible Using a Resorbable Metal Magnesium Membrane and Fixation Screws: A Case Report, Case Rep. Dent. 2024 (2024) 2659893. https://doi.org/10.1155/crid/2659893.

[94]	S.E.C. Pichardo, F.W. Ten Broek, J.P. Richard Van Merkesteyn, Treatment of pathologic fractures of the mandible in stage III medication-related osteonecrosis of the jaw—an observational study, J. Cranio-Maxillofac. Surg. 46 (2018) 1241–1246. https://doi.org/10.1016/j.jcms.2018.05.025.

[95]	A. Neumann, K. Kevenhoerster, Biomaterials for craniofacial reconstruction, GMS Curr. Top. Otorhinolaryngol. Head Neck Surg. 8 (2009) 1–17. https://doi.org/10.3205/cto000060.

[96]	M.E. Pontell, E.B. Niklinska, S.A. Braun, N. Jaeger, K.J. Kelly, M.S. Golinko, Resorbable Versus Titanium Rigid Fixation for Pediatric Mandibular Fractures: A Systematic Review, Institutional Experience and Comparative Analysis, Craniomaxillofacial Trauma Amp Reconstr. 15 (2022) 189–200. https://doi.org/10.1177/19433875211022573.

[97]	Ž.P. Kačarević, P. Rider, A. Elad, D. Tadic, D. Rothamel, G. Sauer, F. Bornert, P. Windisch, D.B. Hangyási, B. Molnar, T. Kämmerer, B. Hesse, E. Bortel, M. Bartosch, F. Witte, Biodegradable magnesium fixation screw for barrier membranes used in guided bone regeneration, Bioact. Mater. 14 (2022) 15–30. https://doi.org/10.1016/j.bioactmat.2021.10.036.

[98]	C. Li, C. Guo, V. Fitzpatrick, A. Ibrahim, M.J. Zwierstra, P. Hanna, A. Lechtig, A. Nazarian, S.J. Lin, D.L. Kaplan, Design of biodegradable, implantable devices towards clinical translation, Nat. Rev. Mater. 5 (2019) 61–81. https://doi.org/10.1038/s41578-019-0150-z.

[99]	L. Sun, Y. Zeng, Z. Shen, C. Yue, Y. Yang, J. Gao, J. Zhang, Q. Yuan, L. Cha, Biodegradable Metal-Based Stents: Advances, Challenges, and Prospects, J. Funct. Biomater. 16 (2025) 315. https://doi.org/10.3390/jfb16090315.

[100]	K. Chen, X. Gu, H. Sun, H. Tang, H. Yang, X. Gong, Y. Fan, Fluid-induced corrosion behavior of degradable zinc for stent application, J. Mater. Sci. Technol. 91 (2021) 134–147. https://doi.org/10.1016/j.jmst.2021.02.050.

[101]	C. Li, Y. Xian, Y. Zhou, X. Ding, J. Sun, D. Xue, Generative Design for Alloys: Harnessing Generative Models for Faster Discovery, Adv. Mater. 38 (2026) e20478. https://doi.org/10.1002/adma.202520478.

[102]	M. Moravej, D. Mantovani, Biodegradable Metals for Cardiovascular Stent Application: Interests and New Opportunities, Int. J. Mol. Sci. 12 (2011) 4250–4270. https://doi.org/10.3390/ijms12074250.

[103]	C. Chen, R. Yue, J. Zhang, H. Huang, J. Niu, G. Yuan, Biodegradable Zn-1.5Cu-1.5Ag alloy with anti-aging ability and strain hardening behavior for cardiovascular stents, Mater. Sci. Eng. C 116 (2020) 111172. https://doi.org/10.1016/j.msec.2020.111172.