Response to Comments

Manuscript: A Mechanistic Framework for Predicting the Biodegradation of Zinc Implants

Authors: Henaka Ariyarathna, Xuan Yang, Aijun Huang, Yunlong Tang, Yuman Zhu

Target journal: Advanced Science

# A. Fix before submission

## A1. The kinetic parameters look inconsistent with the results.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| A1.1 | With k_ORR = 0.51 mm h⁻¹ and C_O₂ = 3.5 mg L⁻¹, Eq. 15 caps recession at about 0.009 mm yr⁻¹. That is about 0.7 µm in 28 days, or about 0.1% disc mass loss. | We thank the reviewer for this dimensional consistency check, and we confirm the estimate. In consistent units, C_O₂ = 3.5 mg L⁻¹ = 1.09 × 10⁻¹⁰ mol mm⁻³ (1 L = 10⁶ mm³) and [Zn]_sol = 7.14/65.38 g cm⁻³ = 1.09 × 10⁻⁴ mol mm⁻³. For k_ORR = 0.51 mm h⁻¹, Eq. 15 then gives v_ORR = 2 k_ORR C_O₂/[Zn]_sol = 1.0 × 10⁻⁶ mm h⁻¹ ≈ 0.009 mm yr⁻¹, i.e. about 0.7 µm in 28 days. The simulated corrosion rates (about 0.12 mm yr⁻¹ for the disc and 0.155 mm yr⁻¹ for the stent) and the measured disc mass loss exceed this value, so the equation as written and the reported results were inconsistent. We have audited the implementation of k_ORR, the regularized interfacial delta function δ_Γ and the oxygen concentration (units, g versus mol, and the interface area scaling) and found [describe finding: e.g. the unit/definition used in the code and the correction made]. [Corrected k_ORR, Table 2, Figures 3–8 and associated text: describe what was re-run and what changed.] | Draft |
| A1.2 | Fig. 5b peaks at about 0.12 mm yr⁻¹, the measured disc loses 0.31%, and the stent peaks at 0.155 mm yr⁻¹. | See response to A1.1. The disc and stent corrosion rates in Figure 5b and Figure 6b are [confirmed / re-computed] after the audit: [values]. | Draft |
| A1.3 | Even k_ORR = 0.1 in Fig. 4c gives 0.28% in 168 h, which is above that ceiling. | See response to A1.1. The mass-loss response for k_ORR = 0.1 in Figure 4c is [confirmed / re-computed] after the audit: [values]. | Draft |
| A1.4 | Ask the student to check the units and definition of k_ORR, δ_Γ and the O₂ concentration, and to show a mass balance (O₂ consumed vs Zn lost). | We checked the units and definitions of k_ORR, δ_Γ and C_O₂ (see A1.1) and added a global mass-balance check, based on the stoichiometry 1 mol O₂ : 2 mol Zn, comparing the O₂ consumed at the interface with the Zn lost over 672 h: [result, to be added to the Supplementary Information as Section S[x]]. | Draft |
| A1.5 | My arithmetic is a back-of-the-envelope check, so they should confirm it. | Confirmed; see A1.1. | Draft |

## A2. The film doesn't seem to cause the slowdown the paper attributes to it.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| A2.1 | The Fig. 5c colour bar peaks at F/F_max ≈ 0.003. With Eq. 12, D_eff is then reduced by only about 0.3%. | We thank the reviewer for this check. If Figure 5c showed F/F_max ≈ 0.003, Eq. 12 would indeed reduce the effective diffusivity by only about 0.3% (D_eff/D ≈ 0.997), so the reviewer's arithmetic is correct. However, the colour bar in Figure 5c was mislabelled: it shows the dimensional corrosion-product concentration F in g mm⁻³, not the normalized quantity F/F_max. The maximum plotted value, F = 0.0029985 g mm⁻³, corresponds to F/F_max = 0.0029985/0.0042045 = 0.713 (F_max = ρ_ZnO(1 − ε) = 5.606 × 0.75 g cm⁻³ = 4.2045 × 10⁻³ g mm⁻³), i.e. about 71% of the maximum corrosion-product concentration. Accordingly, Eq. 12 gives D_eff/D ≈ [0.376 with ε/τ = 0.125, as in Eq. 12 and Table 1; 0.332 only if ε/τ² is used — confirm the tortuosity definition implemented in the code], a reduction in diffusivity of approximately [62% / 67%] at the location of maximum F, rather than 0.3%. We have corrected the colour-bar label and revised the figure caption and associated discussion to distinguish F from F/F_max. | Draft |
| A2.2 | The k_f sweep (Fig. 4a) changes mass loss by only about ±6% for a 4× change in k_f. | We agree that the mass-loss response to variations in k_f over the tested range (50–200) is comparatively modest, consistent with the global sensitivity analysis (Figure 3a), which ranks k_f below C_O₂,0, k_ORR, k_d and D_O₂, and with the stronger response to k_ORR in Figure 4. This does not, however, imply that the corrosion-product field is unimportant: in the ablation simulations described in the response to A2.4/A2.5, removing the field entirely (D_eff = D, or k_f = 0) increases the 672 h mass loss from 0.33% to 0.74% (2.2×). The weak dependence on k_f within 50–200 therefore reflects saturation of the corrosion-product effect over this range, not its absence. | Draft |
| A2.3 | The "film" in Fig. 5c extends about 0.5 mm into the electrolyte. It is a diffuse precipitate in solution, not a surface layer. Real films are micrometres thick. | We agree. In the model, the corrosion-product field F (Eq. 11) is evolved throughout the electrolyte wherever dissolved Zn²⁺ is present and is not confined to the metal surface, so the region of non-zero F in Figure 5c (extending about 0.5 mm into the electrolyte) is a diffuse, volumetric corrosion-product accumulation, not a physically resolved, micrometre-thick surface film. We have revised the manuscript accordingly: (i) in Results (Model Validation and Mechanistic Interpretation), "a distinct film layer", "a dense near-surface region" and "the film thickened" now read "a diffuse corrosion-product accumulation", "a homogenized corrosion-product region" and "the near-surface corrosion-product field developed"; (ii) in Methods, we added a sentence stating that F is a homogenized, volumetric concentration field that is not confined to the metal surface and reduces the local diffusivity through the effective-diffusivity relation (Eq. 12); (iii) "film formation/dissolution/coverage" has been replaced by "corrosion-product formation/dissolution/coverage" in the Methods, Results (Figure 4 discussion and caption) and the captions of Figures 1 and 2; and (iv) the Discussion now states that the corrosion products are treated as a homogenized ZnO-based field rather than a discrete surface film. The Figure 5c colour-bar label has also been corrected (see response to A2.1). | Done |
| A2.4 | The √t-like decay and the O₂ depletion over about 1 mm in Fig. 5d look like transient diffusion in an unstirred electrolyte. | We agree that part of the deceleration can arise from transient oxygen diffusion in an unstirred electrolyte, and that the original simulations did not separate this from additional resistance due to the corrosion-product field. To separate the two, we performed ablation simulations of the disc (672 h, same numerical settings; ratios between runs are computed on the same coarse mesh, and absolute values are scaled to the fine-mesh disc by a single constant mesh factor, C = 1.61, obtained from the 672 h baseline): (R0) baseline; (R1) D_eff = D, i.e. no effect of corrosion products on diffusivity; (R1b) k_f = 0, i.e. no corrosion-product formation; and domain-size / boundary tests (R2, R3; see response to A3). With the corrosion-product effect removed, the mass loss still decelerates (effective time exponent of mass loss over 100–672 h n = 0.69, versus n = 0.40 for the baseline), confirming that oxygen depletion in the static electrolyte contributes to the slowdown, as the reviewer suggests. However, the deceleration is much stronger with the corrosion-product field: the mass loss without the field is 1.06, 1.21, 1.56, 1.74 and 2.2 times that of the baseline at 24, 72, 168, 336 and 672 h, respectively (0.744% against 0.333% at 672 h), so the contribution of the corrosion-product field is negligible at early times and grows progressively. The results are shown in Figure S[x] and Table S[x]. | Draft |
| A2.5 | The "self-limiting corrosion-product mechanism" in the abstract and conclusions is therefore not demonstrated. The student needs to show what causes the deceleration. Options: switch the film off, or set D_eff = D. | We have revised the manuscript to reflect these results. The ablation shows that the deceleration has two contributions: oxygen depletion in the static electrolyte, which persists when corrosion products are switched off, and progressive accumulation of corrosion products, which reduces the effective oxygen diffusivity and lowers the 672 h mass loss by 55% (0.333% against 0.744% for D_eff = D). Setting D_eff = D and k_f = 0 give almost the same result (0.744% and 0.730%), indicating that the corrosion products act through the reduction of diffusivity rather than through Zn²⁺ consumption. We have therefore replaced the statement of a self-limiting corrosion-product mechanism in the abstract, Results and Conclusions by [revised wording, e.g. "degradation slows owing to oxygen depletion in the static electrolyte, which is amplified by progressive corrosion-product accumulation that restricts oxygen transport"], and added the ablation results as Figure S[x] and Table S[x]. | Draft |
| A2.6 | They should also explain what "film coverage (%)" in Fig. 4 means. | We agree that "film coverage (%)" was ambiguous. The quantity plotted in Figure 4d–f is the normalized corrosion-product saturation F/F_max (%), where F is the dimensional corrosion-product concentration (g mm⁻³) and F_max = ρ_ZnO(1 − ε) = 4.2045 × 10⁻³ g mm⁻³; it is evaluated as [state the spatial average, e.g. averaged over the metal–electrolyte interface / over the electrolyte domain / at a specified point]. We have renamed the quantity in the y-axis labels of Figure 4d–f, in the Figure 4 caption and in the text (Methods and the Figure 4 discussion) to "corrosion-product saturation, F/F_max", and the definition is now stated in [Methods / caption]. | Draft |

## A3. The experiment and the model don't match.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| A3.1 | The medium was refreshed every 48 h, which resets O₂ and Zn²⁺. The model seems to have no refresh and no natural convection. | We agree and have clarified this in the manuscript. The simulations do not include the 48 h medium renewal or natural convection (u = 0). Dissolved oxygen is supplied by continuous replenishment at the electrolyte–air interface, which stands in for oxygen resupply in the experiments. This is now stated explicitly in Methods (Numerical Implementation and Simulation Setup): "The 48 h medium renewal used in the immersion tests was not modeled explicitly; instead, dissolved oxygen was assumed to be continuously replenished at the electrolyte–air interface, no natural convection was considered (static conditions), and Zn²⁺ accumulation was not reset." To quantify the effect of the oxygen supply condition, we performed additional simulations with a fixed oxygen concentration at the domain boundary (Figure S[x]), which give a 672 h mass loss about 1.3 times that of the baseline and are independent of domain size (see response to A3.2). We note that Zn²⁺ accumulation is not reset in the model, which affects the Zn²⁺ comparison in Figure 5e [add comment as appropriate]. | Draft |
| A3.2 | The model's electrolyte domain is at least 3× the implant size, and the disc and stent therefore have different domains. The O₂ path to the air interface differs, which confounds the delayed diffusion transition (100 h vs 40 h) attributed to geometry. | We clarify that the disc and the stent were simulated in the same electrolyte domain (30 × 30 × 30 mm; Table S4), with identical boundary conditions, including continuous oxygen replenishment at the electrolyte–air interface. The statement "at least three times larger than the implant characteristic dimension" described the minimum size criterion (the 10 mm disc is the limiting case), not different domain sizes for the two geometries. We have revised the Methods to state the common domain explicitly: "…embedded within a 30 × 30 × 30 mm electrolyte domain, identical for the disc and the stent and larger than three times the implant characteristic dimension in each direction, to minimize boundary effects." The difference in transition time between the disc (about 40 h) and the stent (about 100 h) therefore arises under identical oxygen supply conditions. We acknowledge that the absolute transition time depends on the domain (see the domain-size and fixed-oxygen tests in the response to A2.4), and we have [qualified the corresponding statements in the Results to describe the stent/disc difference as a relative one / added the stent simulation in the fixed-oxygen domain test]. | Draft |
| A3.3 | C_O₂ = 3.5 mg L⁻¹ is about half of air-saturated saline at 37 °C (about 6–7 mg L⁻¹). C_O₂ was the most sensitive parameter in the Morris analysis, yet it was fixed. k_ORR was then fitted, so it may absorb the difference. | [To be completed] | Open |
| A3.4 | Ask for: the vessel geometry and liquid height in the model; a domain-size and refresh sensitivity test; a justification for 3.5 mg L⁻¹, applied separately to in vitro and in vivo. | [To be completed] | Open |

## A4. "Independent validation" is weaker than it sounds.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| A4.1 | Calibration used Liu's r-SBF data (Tris-buffered, 14 days). Validation used HBSS data (28 days), and the two media differ chemically. | [To be completed] | Open |
| A4.2 | The model has a single ZnO phase and no phosphate or carbonate. The paper justifies this with in vivo data showing that complex phases appear only after about 4.5 months. In vitro media with phosphate and carbonate form these phases within days, as the authors' own refs 42 and 44 show. | [To be completed] | Open |
| A4.3 | The validation RMSE (0.0159%) is lower than the calibration RMSE (0.042%), which is unusual and needs explaining. | [To be completed] | Open |
| A4.4 | The fit sits well inside the experimental noise: 28-day SD is 0.115%, n = 3, 5 points. Report normalised error (about 5–8% of the mean), not just RMSE and R². | [To be completed] | Open |
| A4.5 | The earliest data point is at 24 h, but the reaction-to-diffusion transition is claimed at about 15–40 h. Nothing constrains that. | [To be completed] | Open |
| A4.6 | The 14 d → 28 d difference (3.03 vs 3.67 mg) is smaller than the SD. | [To be completed] | Open |
| A4.7 | A Pearson r of 0.95 for two monotonic saturating curves is not a meaningful validation of Zn²⁺ and pH. The model overpredicts Zn²⁺ by about 45% at 336 h. | [To be completed] | Open |

## A5. Calibration identifiability is not shown.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| A5.1 | Fig. 3b: the RMSE of 0.046% was already found at iteration 6, versus 0.042% at "the optimum" (iteration 22). The plateau at about 0.058% suggests a flat landscape. | [To be completed] | Open |
| A5.2 | k_f is the least influential of the calibrated parameters in Fig. 3a. Ask for parameter correlation, a profile likelihood or a posterior. | The relative influence of k_f is addressed in the response to A2.2 (Figure 3a ranks k_f below C_O₂,0, k_ORR, k_d and D_O₂). [Parameter correlation / profile-likelihood analysis to be completed.] | In progress |
| A5.3 | The GSA excluded ε, τ, F_max and [Zn]_sat, which control the film mechanism, and took its parameter ranges from unstated sources. | [To be completed] | Open |
| A5.4 | "Sensitivity to O₂ matches experiments, so the model reproduces oxygen-controlled behaviour" is circular. The model assumes ORR control. | [To be completed] | Open |

## A6. The stent result needs a baseline.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| A6.1 | The paper says mass loss scales about linearly with SA:V. A disc-calibrated rate scaled by SA:V may give a similar or higher number (rough estimate 8–9%). Without that comparison, the 6% doesn't show the mechanistic model adds anything. | [To be completed] | Open |
| A6.2 | Show the baseline. The real added value is the spatial distribution. | [To be completed] | Open |
| A6.3 | The stent dimensions are not given in the main text. "Representative of designs in [43]" cites a review, not the stent in ref 10, which supplied the in vivo number. | [To be completed] | Open |
| A6.4 | In vivo there is flow, protein and different O₂ supply. The abstract and conclusions should say "same order of magnitude" rather than "consistent with in vivo". | [To be completed] | Open |

# B. Should fix

## B1. Model formulation.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| B1.1 | The min() in Eq. 17 looks redundant with the Robin O₂ boundary condition, and it can break the O₂–Zn mass balance. | [To be completed] | Open |
| B1.2 | [Zn]_sat is used in Eq. 14 but not as a precipitation threshold in Eq. 11. Zn²⁺ reaches 57 mg L⁻¹, well above the 10 mg L⁻¹ value used. | [To be completed] | Open |
| B1.3 | F_max is molar, but k_d's units (mm⁶ g⁻² h⁻¹) imply a mass basis. | [To be completed] | Open |
| B1.4 | Explain why Cl⁻ enters as C_Cl² when Eq. 5 is fourth-order, and why Cl⁻ is "catalytic" when Eq. 5 consumes it. | [To be completed] | Open |
| B1.5 | Say how the film is placed relative to the receding interface (the ZnO/Zn volume ratio is about 1.6). | [To be completed] | Open |
| B1.6 | Say how δ_Γ is regularised. | [To be completed] | Open |

## B2. Numerics.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| B2.1 | 18.9M elements for the disc is hard to square with "low-dimensional" and "computationally tractable", especially when the paper says transport is essentially 1D. State the cost, and consider an axisymmetric model for calibration. | [To be completed] | Open |
| B2.2 | Give the interface mesh size relative to the 3–8 µm recession. The jagged surfaces in Fig. 7a look like discretisation noise. | [To be completed] | Open |

## B3. Fig. 4.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| B3.1 | The sweeps use k_f = 100, k_d = 30 and k_ORR = 1.0 as the baseline, not the optimised Table 2 values. The baseline gives 0.39% at 168 h versus about 0.19% in the calibrated Fig. 5a. | [To be completed] | Open |

## B4. Wire study.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| B4.1 | The text gives 2.1–2.7 µm for both the curved and the spaced wires. Fig. 8c shows about 2.4–2.7 µm for curvature and about 2.1–2.4 µm for spacing. | [To be completed] | Open |
| B4.2 | Define "±1 std" in Fig. 8c. | [To be completed] | Open |
| B4.3 | Explain the 13.2× ratio for the 150° arc, given that the stent crown in Fig. 7 shows only about 1.8× (8.55/4.76). The ratio is probably a very small denominator. | [To be completed] | Open |

## B5. Experimental detail.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| B5.1 | Give the balance resolution and whether cleaning blank controls were run. Masses are about 0.5 mg on a 1.1 g disc. | [To be completed] | Open |
| B5.2 | Give the chromic acid temperature and time, disc orientation, CO₂/incubator conditions and pH of HBSS. | [To be completed] | Open |
| B5.3 | Measure Zn²⁺ by ICP in the refreshed solution as an independent check on mass loss. | [To be completed] | Open |
| B5.4 | Ideally add one mechanism test, such as film thickness from FIB or SEM cross-sections, or a stirred or different-O₂ experiment. | [To be completed] | Open |

## B6. Positioning.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| B6.1 | Give a clearer comparison with Chen et al. [23] and the Mg level-set models [18,19]. "Predicting" and "transferable" in the title and Intro are strong for one geometry and one medium. | [To be completed] | Open |

# C. Presentation and consistency

## C1. Availability statements.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| C1.1 | The Code Availability statement says GitHub but links a Zenodo DOI. There is no Data Availability Statement, which Wiley requires. | [To be completed] | Open |

## C2. Wording consistency.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| C2.1 | The Results call the framework "computationally tractable" but the Discussion says it is "computationally demanding". | [To be completed] | Open |
| C2.2 | The text cites "Section 2.3" but the headings are not numbered. | [To be completed] | Open |
| C2.3 | Fix "a R²", "literature" capitalised mid-sentence, and "shown in (Figure 4g–i)". | [To be completed] | Open |

## C3. References.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| C3.1 | Ref 14 and ref 25 spell AlBaraghtheh differently. | This has been published with the same different names. The difference comes from the capitalization of the B. Since it will not be updated | Closed |
| C3.2 | Refs 34 and 35 are two CRC editions (1991 and 2017). | Updated both references to the latest edition | Closed |
| C3.3 | Ref 31 has a duplicated title. | Updated | Closed |
| C3.4 | Re-verify the 2026 references (9, 23, 30). I couldn't check article numbers beyond my knowledge. | Rechecked | Closed |

## C4. Definitions.

| ID | Comment | Response/action taken (with page, line, figure or table) | Status |
| --- | --- | --- | --- |
| C4.1 | Define the Damköhler number and the "transition" criterion in the main text. | [To be completed] | Open |

# Additional changes made by the authors

[List any changes not requested by the reviewer.]