Computational Design of Biodegradable Metallic Implants: Modeling, Optimization, and Clinical Translation

Henaka Ariyarathna​a,b, Aijun Huang​a,b,c, Yunlong Tang​a,b,c*, Yuman Zhu​a,c*

​aMonash Centre for Additive Manufacturing, Monash University, Wellington Rd, Clayton, VIC 3800, Australia

​bDepartment of Materials Science and Engineering, Monash University, Wellington Rd, Clayton, VIC 3800, Australia

​cDepartment of Mechanical and Aerospace Engineering, Monash University, Wellington Rd, Clayton, VIC 3800, Australia

Abstract

Zinc has emerged as a promising biodegradable metal for temporary implant applications. Despite extensive experimental studies, predictive computational frameworks for Zn biodegradation under physiological conditions remain limited. In this study, a three-dimensional mechanistic reaction-diffusion framework is developed by coupling oxygen reduction reaction (ORR) kinetics, mass transport, corrosion-product evolution, and moving-interface tracking. Initial sensitivity analysis showed that dissolved oxygen availability as the primary control on Zn biodegradation. The model was calibrated against literature-reported immersion mass-loss data and predicted 28-day immersion experiments of pure Zn in Hanks' Balanced Salt Solution, yielding a root-mean-square error of 0.0159%. Application of the calibrated model to a six-crown Zn stent, without parameter refitting, produced degradation trends consistent with reported in vivo observations and demonstrated that geometry-dependent material loss emerges naturally from local variations in oxygen transport and corrosion-product evolution. Overall, the proposed framework establishes a predictive foundation for the design and evaluation of biodegradable Zn implants.

Keywords: Biodegradable implant; Zinc; Corrosion modeling; Oxygen reduction kinetics; Ionic transport

# Introduction

Modern medicine increasingly relies on metallic implants to restore physiological function and provide structural support in orthopedic, cardiovascular, and reconstructive applications. Metals are particularly suited to these applications because their high strength, fracture toughness, fatigue resistance, and load-bearing capacity generally exceed those of polymers and ceramics [1]. Conventional implant materials, including stainless steels, titanium alloys, and cobalt–chromium alloys, were developed to resist corrosion and maintain long-term mechanical stability in physiological environments [2–5]. Although this durability has enabled their widespread clinical use, permanent retention can be undesirable when an implant is required only during a limited period of tissue healing. Long-term implant presence may contribute to stress shielding, chronic local reactions associated with wear or metal-ion release, interference with tissue growth and remodeling, and vascular complications such as thrombosis or restenosis. Moreover, temporary fixation devices may require surgical removal after healing, increasing the risks of anesthesia, infection, additional tissue damage, prolonged recovery, and healthcare costs [6–10]. Biodegradable metallic implants have therefore emerged as an alternative strategy to provide mechanical support during the critical healing period and then degrade into products the body can tolerate, metabolize, or clear [11–15]. Their gradual degradation may eliminate the need for removal surgery while progressively transferring mechanical load from the implant to the recovering tissue. Magnesium (Mg), iron (Fe), and zinc (Zn), together with their alloys, are the principal metallic systems investigated for biodegradable implant applications because they combine clinically relevant mechanical properties with the capacity to degrade under physiological conditions [16–20]. Researchers are exploring these materials for orthopedic fixation devices, vascular scaffolds, porous bone implants, and other temporary biomedical devices. However, their degradation behavior and functional performance depend strongly on material composition, microstructure, surface condition, implant geometry, manufacturing route, mechanical loading, and the surrounding biological environment. Developing biodegradable metallic implants therefore requires an integrated design approach that provides sufficient mechanical support during healing while controlling the rate, morphology, and biological consequences of degradation.

Unlike permanent devices, biodegradable metallic implants are dynamic systems whose geometry, surface condition, mechanical capacity, and biological interactions evolve continuously after implantation. Coupled electrochemical reactions, mass transport, corrosion-product formation, mechanical loading, and local tissue responses govern their degradation, and material composition, microstructure, manufacturing history, and implant geometry further influence it [21–23]. Consequently, degradation measured using simplified laboratory specimens may not represent the behavior of functional screws, plates, scaffolds, or stents in vivo [21,24]. Importantly, average mass loss alone cannot define implant performance. Uniform surface recession and localized corrosion can produce similar total mass loss but markedly different residual strength, particularly when degradation occurs at critical features such as screw threads, plate holes, scaffold junctions, or stent crowns [25–28]. A successful implant must therefore retain adequate mechanical function during the critical healing period while gradually transferring load to the recovering tissue and degrading without harmful accumulation of corrosion products. Biodegradable implants should therefore be designed around their complete degradation–mechanical performance trajectory, rather than their initial properties or nominal corrosion rate alone. This time-dependent behavior requires an integrated design framework that considers degradation, mechanics, transport, manufacturing, and tissue healing as interacting components of a single material–device system [22,24,29].

The design of biodegradable metallic implants is governed by four interdependent mismatches. First, implant degradation must be synchronized with tissue healing: premature degradation may cause loss of fixation or structural support, whereas prolonged persistence may negate the advantages of a temporary device [30,31]. Second, the progressive reduction in implant strength and stiffness must be balanced by the recovery of the surrounding tissue to ensure gradual load transfer without excessive deformation or sudden instability [23,32,33]. Third, ion and corrosion-product release must remain compatible with transport, biological tolerance, metabolism, and clearance; otherwise, local accumulation may occur even when the average degradation rate appears acceptable [31,34,35]. Finally, computational models must balance physical fidelity with experimental and computational feasibility. Although detailed multiphysics models can represent electrochemical reactions, mass transport, film evolution, mechanics, and tissue response, their parameter requirements and computational cost can limit calibration, uncertainty analysis, and design optimization [36–38]. Effective implant design therefore requires models that capture the mechanisms controlling clinically relevant outcomes, remain identifiable from available experiments, and are efficient enough for device-scale simulation.

This review critically examines the computational design of biodegradable metallic implants across Mg-, Fe-, and Zn-based systems, emphasizing the relationships among material properties, degradation behavior, evolving mechanical performance, implant geometry, manufacturing, and tissue healing. It covers empirical, phenomenological, stochastic, and mechanistic degradation models; degradation-coupled structural and mechanobiological simulations; porous and topology-optimized architectures; additive-manufacturing-aware design; surrogate modeling; machine learning; uncertainty quantification; and model verification and validation [32–38]. Rather than organizing these developments into separate material- or method-specific topics, the review introduces a unified, lifetime-aware computational design framework that evaluates implants over their full functional period. Within this framework, material selection and implant geometry determine the initial response; degradation models predict the evolution of morphology and material loss; mechanical models quantify the resulting decline in device function; and tissue-healing models describe the progressive recovery of biological load-bearing capacity. Optimization then identifies designs that maintain application-specific performance throughout healing while satisfying degradation, biological, manufacturing, and safety constraints. The principal novelty of this review is integrating previously fragmented computational approaches into a common pathway linking material-scale degradation mechanisms to device-scale performance and clinical translation. This review pays particular attention to model transferability among Mg, Fe, and Zn systems; the identification and calibration of degradation parameters; the limitations of extrapolating from coupons to functional devices; and the experimental and uncertainty evidence needed for computational predictions to support clinically relevant design decisions.

This review follows the complete computational design pathway for biodegradable metallic implants. It first defines application-specific clinical requirements and the time-dependent relationship between implant degradation, mechanical support, and tissue healing. The review then compares Mg-, Fe-, and Zn-based systems from a computational perspective and examines empirical, phenomenological, stochastic, and mechanistic approaches to modeling degradation. Subsequent sections address the coupling of degradation with structural mechanics, implant–tissue interactions, and evolving functional performance, followed by computational methods for geometry optimization, porous and graded architecture design, and additive-manufacturing-aware development. The review then evaluates the roles of sensitivity analysis, surrogate modeling, machine learning, uncertainty quantification, and multi-objective optimization. Finally, the review considers model calibration, verification and validation, experimental-to-device translation, regulatory credibility, and the remaining barriers to clinical implementation. This sequence connects fundamental degradation physics with implant-level design decisions and provides a structured pathway toward predictive, lifetime-aware, and clinically translatable computational design of biodegradable metallic implants.

# Clinical design requirements

## Application classes

Table 1 summarizes the strength, stiffness, device-specific, degradation and material requirements for the device classes discussed below.

Table 1. Application-specific requirements for biodegradable metallic implants (UTS: ultimate tensile strength; YS: yield strength; SS: stainless steel).

| Device class | Strength and stiffness | Device-specific performance | Degradation and support period | Material trends (Mg / Zn / Fe) |
| --- | --- | --- | --- | --- |
| Orthopedic fixation (screws, pins, plates, nails) | UTS >~300 MPa, YS >~200 MPa; modulus closer to cortical bone (10–30 GPa) than permanent alloys, to limit stress shielding [41,42] | Screws: pull-out ~200–800 N (Mg), ~3800 N (Fe–SS hybrid) [47,48]. Pins and K-wires: Mg bending strength 437–625 MPa [49,50]. Plates and nails: four-point bending and torsional fatigue testing [51,52] | LAE442 Mg pins lose ~40% of bending strength within 9 months and retain ~12% after 3.5 years [51,54] | Moduli: Mg 41–45 GPa [44], Zn 90–110 GPa [45], Fe 190–211 GPa [46]; Fe and Zn reach 400–900 MPa, pure Mg lower [51,52] |
| Porous scaffolds | Porosity 50–90%; porosity increases reduce stiffness and strength and raise strut stress concentrations [58,59,61] | Pore size ~100–600 μm (300–600 μm favors vascularization); interconnectivity >30% with windows >100 μm [58–60] | Architecture governs permeability, transport and local corrosion environment, and so degradation rate and uniformity [16,62] | Mg: closest modulus to trabecular bone but fastest degradation; Zn: slower, controlled, topology-sensitive; Fe: strongest but slow, needs high porosity [64,65] |
| Vascular stents | YS >200 MPa, UTS >300 MPa; Zn–Cu stents reach radial strength ~120 kPa [41,68,69] | Recoil <~4% [41,67]; fatigue life 10–20 million cycles [41,68]; strut thickness: PLLA 150–170 μm vs thinner metallic struts [71] | Integrity for ~3–6 months, absorption over 12–24 months (~0.02 mm yr⁻¹) [41,68]; corrosion raises recoil gradually [28] | Zn: within the therapeutic window, no hydrogen evolution (unlike Mg) [41,69] |

### Orthopedic Fixation Devices

Orthopedic fixation devices, including screws, pins, Kirschner wires, plates, and intramedullary nails, are intended to provide immediate mechanical stabilization after implantation, maintain structural integrity during the critical early stages of fracture repair, and progressively transfer load to the regenerating bone as healing advances [39,40]. Although their geometries and loading conditions differ considerably, these devices share a common biomechanical requirement: they must provide sufficient strength for temporary load bearing while minimizing mechanical mismatch with the surrounding skeletal tissue. Accordingly, biodegradable fixation systems are generally expected to achieve tensile and yield strengths exceeding approximately 300 MPa and 200 MPa, respectively, while maintaining elastic moduli closer to cortical bone (10-30 GPa) than conventional permanent implant materials [41,42]. This stiffness compatibility is particularly important because excessive implant rigidity can induce stress shielding, whereby the implant bears a disproportionate share of physiological loading, reducing mechanical stimulation of the adjacent bone and impairing remodeling and healing [40,43]. Among biodegradable metals, magnesium alloys exhibit the most favorable stiffness profile, with elastic moduli of approximately 41-45 GPa [44], whereas zinc alloys typically possess intermediate moduli of 90-110 GPa [45], and iron-based alloys exhibit substantially higher values of 190-211 GPa, making them more susceptible to stress-shielding effects despite their superior mechanical strength [46].

Within this shared framework, geometric design fundamentally influences the mechanical demands placed on each fixation device. Screws and interference screws are primarily evaluated through pull-out strength, insertion torque, and torsional resistance because their clinical performance depends on achieving sufficient anchorage within bone and resisting loosening under physiological loading before biological fixation develops. Reported pull-out strengths vary widely depending on material and design, ranging from approximately 200 to 800 N for magnesium-based screws and reaching nearly 3800 N in friction-welded iron-stainless steel hybrid systems developed to enhance fixation strength [47,48]. In contrast, pins and Kirschner wires have smaller cross-sectional dimensions and therefore lower structural reserves, making them particularly sensitive to the balance between strength and ductility. Magnesium alloy pins exhibit bending strengths of 437-625 MPa, substantially exceeding the 224-271 MPa reported for self-reinforced PLLA systems, highlighting the mechanical advantage provided by biodegradable metals in minimally invasive fixation applications [49,50]. Plates and intramedullary nails operate under considerably higher absolute loading conditions and are consequently assessed using four-point bending and torsional fatigue testing rather than fixation strength alone. Although iron- and zinc-based systems can achieve tensile strengths exceeding 400-900 MPa, pure Mg constructs generally exhibit lower load-bearing capacity and often require alloy strengthening or hybrid structural designs to meet the mechanical demands of load-bearing orthopedic applications [51,52]. For example, LAE442 intramedullary pins lose approximately 40% of their bending strength within nine months in vivo and retain only about 12% of their original strength after 3.5 years, motivating extensive research into alloying and surface-modification strategies that can moderate degradation while preserving magnesium's favorable stiffness characteristics [51,54].

### Porous Scaffolds and Bone Defect Substitutes

Porous scaffolds perform a fundamentally different function from solid fixation devices. Rather than providing standalone structural support, they are designed to share loads with surrounding tissue while serving as temporary templates for bone regeneration. Their interconnected porous architecture enables the transport of cells, nutrients, metabolic waste, and vascular networks throughout the construct, thereby supporting tissue infiltration and new bone formation [57]. To satisfy these combined mechanical and biological requirements, scaffold designs typically target total porosities of 50-90%, with pore sizes ranging from approximately 100 to 600 μm. Macropores within this range promote osteoblast migration and direct osteogenesis, while high pore interconnectivity, generally exceeding 30% with interconnecting windows larger than 100 μm, is essential for enabling tissue penetration throughout the scaffold rather than restricting bone growth to the outer surface [58,59]. Larger pores, particularly in the 300-600 μm range, are additionally associated with enhanced vascularization because they provide sufficient space for endothelial invasion and capillary development, which are critical for sustaining viable tissue growth within the scaffold interior [60].

These biological advantages come with significant mechanical penalties. Increasing porosity improves cellular infiltration and mass transport but simultaneously reduces stiffness and strength while amplifying stress concentrations within individual struts, making scaffold design an inherent compromise between biological performance and structural integrity rather than a straightforward optimization of either objective alone [61]. The widespread adoption of additive manufacturing has transformed this design paradigm by enabling local control over scaffold architecture. As a result, unit-cell topology increasingly governs mechanical behavior and degradation kinetics rather than porosity alone. Scaffold architecture strongly influences both mechanical behavior and degradation by affecting permeability, fluid transport, surface area, and local corrosion environments [16,62]. Consequently, scaffold performance depends not only on porosity but also on the interaction between architecture and alloy chemistry.

Material selection therefore remains a critical determinant of scaffold behavior, and the relative characteristics of magnesium-, zinc-, and iron-based systems broadly mirror those observed in solid fixation devices. Magnesium scaffolds offer the closest elastic modulus match to trabecular bone but degrade most rapidly, increasing the risk of premature structural collapse before sufficient bone ingrowth has occurred. Zinc scaffolds generally degrade more slowly and in a more controlled manner, though at the expense of higher stiffness, and they also show pronounced sensitivity to architectural design, as topology can strongly influence both degradation rate and uniformity. Iron scaffolds provide the greatest mechanical robustness but have intrinsically slow corrosion kinetics, often requiring deliberate introduction of high porosity to achieve clinically relevant resorption times [64,65]. This is particularly evident in functionally graded and highly porous iron structures, where increased surface-area-to-volume ratio is intentionally used to accelerate degradation and compensate for the limited corrosion of bulk iron. However, the additional architectural degrees of freedom available in porous structures enable degradation behavior to be tailored through the combined effects of alloy chemistry and topological design, providing a powerful framework for balancing mechanical support, biological integration, and implant resorption [66].

### Cardiovascular Stents

Cardiovascular stents follow a design philosophy similar to orthopedic fixation devices, but under substantially stricter geometric and mechanical constraints. Following deployment, a bioresorbable stent must generate sufficient radial force to maintain vessel patency while minimizing acute elastic recoil, which is generally expected to remain below ~4% to prevent the artery from returning toward its pre-treatment diameter [41,67]. These requirements translate into a demanding combination of mechanical strength, fatigue resistance, and controlled degradation, with bioresorbable vascular scaffolds typically targeting yield strengths above 200 MPa, ultimate tensile strengths exceeding 300 MPa, and fatigue lives of 10-20 million cycles to withstand continuous pulsatile loading. At the same time, mechanical integrity should be retained for approximately 3-6 months before gradual bioabsorption over 12-24 months, corresponding to corrosion rates on the order of 0.02 mm year⁻¹ [41,68]. Among biodegradable metals, zinc alloys have emerged as particularly promising candidates because their corrosion behavior falls within this therapeutic window while providing sufficient radial support. For example, Zn-Cu stent systems have achieved radial strengths of approximately 120 kPa, comparable to commercially available permanent and bioresorbable stents at similar strut dimensions, while avoiding the hydrogen evolution associated with magnesium degradation [41,69].

Beyond material selection, strut geometry strongly influences clinical performance because stents interact directly with the local hemodynamic environment. Increasing strut thickness improves radial stiffness and load-bearing capacity but also protrudes further into the vessel lumen, disturbing endothelial shear stress and promoting flow recirculation zones associated with thrombosis and restenosis [70]. Conversely, thinner struts better preserve physiological blood flow but require higher material strength to maintain equivalent radial support. This trade-off highlights a key advantage of metallic bioresorbable systems over polymeric counterparts. Owing to their substantially lower elastic modulus, PLLA-based scaffolds require strut thicknesses of approximately 150-170 μm, which can be more than twice those of metallic stents providing similar radial strength. The resulting hemodynamic disturbance is now recognized as a major contributor to the elevated thrombosis rates observed in first-generation polymeric bioresorbable scaffolds [71]. Metallic stents therefore offer the potential to combine thinner struts with adequate radial support. Computational studies further demonstrate that progressive corrosion primarily manifests as a gradual increase in radial recoil rather than the abrupt loss of mechanical integrity characteristic of polymer scaffold fracture [28]. These performance objectives are commonly evaluated using standardized protocols, including ASTM F2477 for pulsatile fatigue durability, ASTM F3067 for radial compression testing, and ISO 25539-2 for endovascular device performance, providing the experimental and regulatory framework against which computational and physical stent designs are validated.

## Time-dependent performance requirements

Permanent implant alloys such as 316L stainless steel, Ti-6Al-4V, and Co-Cr were developed under a static design paradigm in which mechanical strength and corrosion resistance are maximized and subsequently maintained throughout the implant lifetime [40]. Biodegradable metals fundamentally depart from this framework. From the moment of implantation, corrosion continuously alters implant geometry, surface condition, and load-bearing capacity, so initial mechanical properties alone offer little indication of long-term functional performance [39,73]. Consequently, biodegradable implants must be evaluated by tracking the temporal evolution of their properties, treating strength, stiffness, and structural integrity as functions of time rather than static material attributes [74].

This evolving behavior directly affects how loads are transferred between the implant and surrounding tissue. In permanent implants, mechanical support remains largely unchanged after implantation, often resulting in prolonged stress shielding as the device continues to bear a substantial fraction of physiological load as healing progresses [40,56]. By contrast, biodegradable implants are intended to progressively relinquish mechanical responsibility as degradation proceeds. As corrosion reduces implant stiffness and effective cross-sectional area, load gradually transfers to the regenerating bone or remodeling vessel wall, enabling restoration of physiological mechanical function [46,74]. The success of this process therefore depends not only on how much support an implant provides, but also on when that support is withdrawn.

Because load transfer must follow biological healing rather than material degradation alone, mechanical retention requirements are inherently time-dependent. Near-complete fixation stability is generally required during the first 4-6 weeks, when inflammatory tissue and soft callus contribute to little intrinsic load-bearing capacity. As healing progresses and hard callus formation begins, fixation devices are typically expected to retain at least 50% of their initial strength in load-bearing applications and approximately 30% in non-load-bearing applications. By the completion of cortical union or vascular remodeling, generally between 6 and 12 months depending on the application, most mechanical responsibility should have shifted to the host tissue [41,53]. Deviations from this schedule in either direction can compromise clinical outcomes: excessively rapid degradation risks fixation failure, vessel recoil, or restenosis before healing is complete, whereas prolonged persistence recreates the stress shielding and chronic foreign-body effects that biodegradable implants are intended to eliminate [75,76].

Achieving this synchronization requires degradation kinetics to fall within a relatively narrow therapeutic window. Magnesium-based implants often degrade too rapidly, with reported corrosion rates of approximately 0.5-2.7 mm year⁻¹, substantially exceeding the rates typically considered compatible with bone healing. Rapid corrosion can lead to premature loss of mechanical support, hydrogen evolution, and localized alkalization [46,77]. Iron-based systems occupy the opposite extreme, often degrading at rates below 0.1 mm year⁻¹ and therefore retaining mechanical function long after healing is complete unless accelerated through porosity engineering or alloying strategies [78,79]. Zinc alloys generally exhibit intermediate degradation rates of 0.1-0.3 mm yr⁻¹, closer to the temporal requirements of many orthopedic and cardiovascular applications, while avoiding the hydrogen evolution associated with magnesium degradation [68].

More broadly, implant performance can be viewed through a healing-degradation performance corridor defined by the host tissue's evolving mechanical competence. The lower boundary of this corridor is the minimum support required to prevent fixation failure, scaffold collapse, or vessel recoil before sufficient tissue regeneration has occurred, whereas the upper boundary reflects the point beyond which persistent implant support becomes detrimental through stress shielding, delayed functional recovery, or prolonged foreign-body presence [40,80]. Importantly, these boundaries are determined by tissue-specific healing kinetics rather than material properties alone, meaning that an identical degradation trajectory may be appropriate in one anatomical setting yet inadequate in another [39,73]. Consequently, maintaining implant performance within this corridor depends not only on alloy chemistry but also on architectural design variables such as porosity, strut thickness, and graded density, which can tailor the temporal evolution of stiffness and degradation to a particular clinical application [74].

These considerations have driven the emergence of lifetime-aware design, in which degradation is treated not as an unavoidable material limitation but as an engineering parameter that can be deliberately tailored to clinical requirements [80]. Within this framework, the objective is no longer to maximize initial strength or minimize corrosion, but to optimize the implant's entire performance trajectory so mechanical support, degradation, and tissue regeneration remain synchronized throughout healing. Consequently, a central challenge for computational design is predicting when implant load-bearing capacity will decline below the level that healing tissue can safely compensate for and using those predictions to engineer patient- and application-specific degradation profiles.

## Failure modes and acceptance criteria

Writing prompt: Cover premature section loss, localized corrosion, fatigue, loss of fixation, excessive persistence, and harmful accumulation of degradation products.

[Add review text and citations here]

Continuous mass loss during biocorrosion can lead to several distinct failure modes in biodegradable Mg-, Zn-, and Fe-based implants, rather than a single outcome governed solely by degradation rate. Implant failure depends not only on how rapidly material is lost but also on the spatial distribution of corrosion, the mechanical loading environment, and the ability of the implant to maintain its intended function throughout healing. Consequently, acceptable performance must be defined through multiple mechanical and biological criteria rather than by a degradation-rate corridor alone [22,40,68].

One of the most important determinants of failure is corrosion morphology. Uniform degradation generally produces a predictable reduction in cross-sectional area and load-bearing capacity, allowing implants to retain functionality until substantial material loss has occurred. In contrast, localized corrosion creates highly detrimental stress concentrations that can trigger structural failure at much lower overall mass losses. Microgalvanic interactions between the matrix and secondary phases or impurities concentrate dissolution into pits that act as crack initiation sites, allowing fracture to occur at less than 5-10% total mass loss [22,81]. Pitting susceptibility is strongly material-dependent. In Mg alloys, intermetallic phases such as Mg₂Ca and Mg₁₇Al₁₂, together with trace impurities such as Fe and Ni, promote localized attack, whereas in zinc alloys the distribution and morphology of secondary phases such as FeZn₁₅ strongly influence whether degradation remains uniform or becomes highly localized [81].

Localized damage becomes even more critical under cyclic physiological loading, where corrosion and fatigue interact synergistically. Repeated loading ruptures protective surface films, continuously exposing fresh metal and transforming corrosion pits into fatigue crack initiation sites. As a result, magnesium alloys can experience fatigue strength reductions exceeding 50% in simulated physiological environments compared with air [41]. The practical consequences are particularly severe for load-bearing devices. Bare LAE442 magnesium intramedullary nails have been reported to fail in fewer than 70,000 bending cycles in Hank's solution, substantially below the approximately 500,000-cycle benchmark typically applied to fracture-fixation devices [82]. Likewise, biodegradable vascular stents must withstand millions of pulsatile loading cycles during their service life, making corrosion-fatigue resistance a critical design requirement [41]. Experimental bioreactor studies further demonstrate that combined fluid perfusion and cyclic loading accelerate pit growth and crack propagation relative to static immersion conditions [55].

In addition to fracture and fatigue failure, degradable fixation devices may fail through progressive loss of fixation stability. Localized corrosion around screw threads, plate holes, and other geometric discontinuities reduces extraction torque, bending stiffness, and overall construct stability before complete structural failure occurs. For example, degrading Herbert screws have exhibited a reduction in bending stiffness from approximately 15 GPa to 8 GPa over 6-12 weeks, whereas magnesium plate-and-screw systems show measurable reductions in fixation strength during early stages of degradation [85,86]. A related geometry-dependent mechanism is stress corrosion cracking (SCC), in which tensile stresses interact with corrosion damage to produce fracture at stresses well below the nominal yield strength. Crevice corrosion occurring in shielded fluid environments, such as plate-bone interfaces and poorly perfused scaffold interiors, can further amplify these localized degradation processes [29,73].

Failure can also arise when degradation is excessively slow. In such cases, implants persist beyond their intended functional lifetime and continue to behave as permanent foreign bodies. This issue is especially evident in iron-based systems, where in vivo studies have demonstrated multi-year structural persistence accompanied by the formation of dense iron oxide and hydroxide layers that inhibit complete resorption [88]. Excessive persistence may prolong stress shielding, maintain chronic foreign-body encapsulation, and, in vascular applications, prevent restoration of natural vasomotion following vessel healing. Long-lived corrosion-product layers may additionally alter load-transfer pathways and fatigue behavior by mechanically shielding the underlying structure from physiological stresses [89].

Beyond mechanical considerations, biodegradable implants must also satisfy biological acceptance criteria related to their degradation products. Magnesium degradation produces hydrogen gas and increases local alkalinity, requiring degradation rates low enough that gas generation remains below tissue clearance capacity. Excessive hydrogen evolution can result in gas-pocket formation, tissue separation, and impaired healing [22]. Zinc and iron do not generate hydrogen gas but instead pose risks from ion release and corrosion-product accumulation. The nature, solubility, and clearance of these degradation products vary considerably between material systems and influence both local tissue responses and systemic safety [15,70]. Consequently, acceptable implant performance requires not only mechanical stability but also degradation-product generation rates that remain within physiological tolerance limits. Taken together, these failure modes demonstrate that successful biodegradable implants must maintain structural integrity throughout healing, resist localized corrosion and corrosion-fatigue damage, preserve fixation stability, avoid excessive persistence, and generate degradation products at physiologically tolerable levels. Acceptance criteria must therefore consider mechanical retention, degradation morphology, biological response, and resorption time simultaneously, rather than relying solely on degradation rate.

## Standards as computational acceptance criteria

Writing prompt: Connect ASTM F3160, ASTM F3268, ISO 10993-15, device-specific mechanical tests, ASME V&V 40, and FDA computational credibility guidance to model inputs and outputs.

[Add review text and citations here]

# Design space of biodegradable metals

## Magnesium systems

Writing prompt: Summarize computationally relevant properties, rapid and localized corrosion, hydrogen evolution, film effects, and environmental sensitivity.

[Add review text and citations here]

## Zinc systems

Writing prompt: Discuss intermediate degradation behavior, oxygen-reduction control, corrosion-product films, localized attack, and the relative shortage of implant-scale mechanistic models.

[Add review text and citations here]

## Iron systems

Writing prompt: Discuss high strength, slow bioresorption, persistent products, and architecture or alloy strategies for accelerating controlled degradation.

[Add review text and citations here]

## Cross-material comparison

Writing prompt: Compare Mg, Zn, and Fe by degradation timescale, localization tendency, mechanical retention, dominant chemistry, and modelling maturity.

[Add review text and citations here]

## Recommended comparison table

Writing prompt: Insert a table comparing material advantage, dominant risk, mechanistic variables, calibration data, and computational design need.

[Add review text and citations here]

# Computational representations of biodegradation

## Empirical geometry-reduction models

Writing prompt: Review uniform surface recession, prescribed thinning, empirical mass-loss laws, and time-dependent geometry updates.

[Add review text and citations here]

## Continuum damage mechanics models

Writing prompt: Review scalar and nonlocal corrosion-damage variables, stiffness degradation, element deletion, mesh dependence, and VUMAT implementations.

[Add review text and citations here]

## Stochastic and cellular models

Writing prompt: Review Monte Carlo, cellular automata, probabilistic surface removal, pit nucleation, growth, and morphology statistics.

[Add review text and citations here]

## Mechanistic reaction-transport models

Writing prompt: Review anodic and cathodic kinetics, species transport, oxygen limitation, pH, buffering, film precipitation and dissolution, chloride effects, and flow.

[Add review text and citations here]

## Interface-tracking methods

Writing prompt: Compare element deletion, level-set, phase-field, ALE moving meshes, immersed methods, and voxel-based updates.

[Add review text and citations here]

## Model hierarchy and selection

Writing prompt: Explain when empirical, phenomenological, reduced mechanistic, or high-fidelity models are appropriate for screening, prediction, optimization, or regulatory evidence.

[Add review text and citations here]

# Coupling degradation with implant mechanics

## Sequential coupling

Writing prompt: Describe degradation prediction followed by geometry or material updating and structural analysis at selected time points.

[Add review text and citations here]

## Fully coupled mechanochemical degradation

Writing prompt: Discuss feedback among stress, strain, corrosion rate, section loss, stress redistribution, cracking, and fatigue.

[Add review text and citations here]

## Bone-implant and vessel-implant interaction

Writing prompt: Cover contact, friction, screw insertion and pretension, fracture-gap mechanics, callus maturation, balloon expansion, plaque contact, recoil, and vessel remodeling.

[Add review text and citations here]

## Residual structural performance metrics

Writing prompt: Compare residual strength, stiffness, fatigue resistance, pull-out force, interfragmentary motion, radial support, recoil, and lumen gain.

[Add review text and citations here]

## Coupling with tissue healing and remodelling

Writing prompt: Review models that jointly represent implant degradation, load transfer, tissue regeneration, bone ingrowth, and biological adaptation.

[Add review text and citations here]

# Geometry, architecture and manufacturing-aware design

## Conventional parameter optimization

Writing prompt: Review optimization of screw threads, core diameter, plate thickness, hole number, stent struts, connectors, and related geometric variables.

[Add review text and citations here]

## Porous and lattice structures

Writing prompt: Discuss porosity, pore size, unit-cell topology, anisotropy, surface-area-to-volume ratio, permeability, tissue ingrowth, and degradation localization.

[Add review text and citations here]

## Topology optimization

Writing prompt: Separate optimization at implantation, robust optimization across degraded states, and true time-dependent topology optimization.

[Add review text and citations here]

## Patient-specific design

Writing prompt: Cover CT segmentation, anatomical reconstruction, virtual implantation, physiological loading, printability, insertion-path checks, and surgeon verification.

[Add review text and citations here]

## Additive-manufacturing constraints

Writing prompt: Discuss minimum feature size, roughness, defects, residual stress, crystallographic texture, dimensional tolerance, and build orientation.

[Add review text and citations here]

## Coatings and spatially graded designs

Writing prompt: Review coating thickness, local surface modification, functionally graded porosity, and variable-stiffness architectures as design variables.

[Add review text and citations here]

# Optimization, surrogate models and uncertainty

## Design variables

Writing prompt: Organize material, kinetic, geometric, coating, architectural, environmental, and patient-specific variables.

[Add review text and citations here]

## Objectives

Writing prompt: Define objectives such as minimizing mass, stress shielding, peak stress, degradation non-uniformity, and harmful ion accumulation while preserving required support.

[Add review text and citations here]

## Constraints and multiobjective formulation

Writing prompt: State time-dependent mechanical, biological, manufacturing, and geometric constraints and explain Pareto trade-offs.

[Add review text and citations here]

## Sensitivity analysis and parameter identifiability

Writing prompt: Discuss local sensitivity, Morris elementary effects, Sobol indices, profile likelihood, parameter correlation, and practical identifiability.

[Add review text and citations here]

## Calibration and Bayesian inference

Writing prompt: Cover deterministic inverse fitting, Bayesian calibration, prior selection, posterior uncertainty, and transfer from coupon to implant geometries.

[Add review text and citations here]

## Surrogates and machine learning

Writing prompt: Review Gaussian processes, polynomial response surfaces, neural surrogates, reduced-order models, physics-informed methods, and active learning.

[Add review text and citations here]

## Robust and reliability-based design

Writing prompt: Discuss uncertainty propagation, probability of failure, worst-case design, reliability constraints, and patient-population variability.

[Add review text and citations here]

# Calibration, validation and model credibility

## Numerical verification

Writing prompt: Address code verification, solution verification, mesh and time-step convergence, conservation checks, and benchmark problems.

[Add review text and citations here]

## Experimental calibration data

Writing prompt: Review mass loss, hydrogen evolution, electrochemistry, surface recession, pH, ion concentration, film composition, and mechanical test data.

[Add review text and citations here]

## Local morphology validation

Writing prompt: Emphasize micro-CT, profilometry, pit-depth distributions, remaining-volume maps, and spatial comparison rather than mass loss alone.

[Add review text and citations here]

## Component-level and in vivo validation

Writing prompt: Discuss screw pull-out, plate bending, scaffold compression, stent radial testing, ex vivo morphology, animal data, and clinical observations.

[Add review text and citations here]

## Uncertainty quantification

Writing prompt: Separate parameter, experimental, model-form, geometric, loading, and biological uncertainties.

[Add review text and citations here]

## Context of use and credibility assessment

Writing prompt: Apply ASME V&V 40 and FDA risk-informed credibility concepts to define the question of interest, model risk, evidence requirements, and decision consequence.

[Add review text and citations here]

# Translational and regulatory considerations

## Role of computation in device development

Writing prompt: Explain how modelling supports design-space exploration, worst-case selection, test planning, interpolation, residual-strength prediction, and evidence integration.

[Add review text and citations here]

## Standards and regulatory pathways

Writing prompt: Discuss absorbable-metal material characterization, degradation testing, biological evaluation, device-specific mechanical standards, and jurisdiction-specific requirements.

[Add review text and citations here]

## Clinical and regulatory case studies

Writing prompt: Use approved or cleared biodegradable metallic devices, such as magnesium fixation screws and vascular scaffolds, to identify expected evidence packages.

[Add review text and citations here]

## Reproducibility and model reporting

Writing prompt: Recommend reporting governing equations, assumptions, parameters, calibration datasets, numerical settings, code availability, uncertainty, and validation limits.

[Add review text and citations here]

## Remaining barriers to translation

Writing prompt: Discuss in vitro-to-in vivo mismatch, limited longitudinal data, patient variability, manufacturing variability, computational cost, and absence of agreed digital endpoints.

[Add review text and citations here]

# Research gaps and future directions

## Lifetime-aware rather than initial-state optimization

Writing prompt: Explain the need to optimize the full support-degradation trajectory rather than only performance at implantation.

[Add review text and citations here]

## Transferable models for Zn and Fe

Writing prompt: Identify the need for material-specific but computationally comparable mechanistic frameworks beyond the Mg-dominated literature.

[Add review text and citations here]

## Coupled degradation-healing digital twins

Writing prompt: Discuss integration of corrosion, tissue regeneration, loading history, imaging, and patient-specific updating.

[Add review text and citations here]

## Multifidelity and data-efficient modelling

Writing prompt: Propose combining high-fidelity mechanistic models, reduced models, experiments, and in vivo data.

[Add review text and citations here]

## Standardized validation datasets and benchmarks

Writing prompt: Recommend open geometries, boundary conditions, degradation datasets, mechanical tests, and reporting conventions.

[Add review text and citations here]

## Regulatory-ready computational evidence

Writing prompt: Outline the path toward predefined contexts of use, credibility plans, traceable software, uncertainty bounds, and decision-relevant validation.

[Add review text and citations here]

# Conclusions

## Main synthesis

Writing prompt: Summarize how degradation physics, structural mechanics, biological response, geometry, and uncertainty must be treated as one design problem.

[Add review text and citations here]

## Design recommendations

Writing prompt: State the most defensible modelling and validation practices for future biodegradable metallic implant studies.

[Add review text and citations here]

## Final perspective

Writing prompt: Conclude with the role of computational design in moving biodegradable metals from material development toward predictable, patient-relevant, and clinically translatable devices.

[Add review text and citations here]

# Suggested figures

[Add and update content here]

# Suggested tables

[Add and update content here]

# References

[1]	M.S. Baltatu, C.A. Tugui, M.C. Perju, M. Benchea, M.C. Spataru, A.V. Sandu, P. Vizureanu, Biocompatible Titanium Alloys used in Medical Applications, Rev. Chim. 70 (2019) 1302–1306. https://doi.org/10.37358/RC.19.4.7114.

[2]	K.-H. Frosch, K.M. Stürmer, Metallic Biomaterials in Skeletal Repair, Eur. J. Trauma 32 (2006) 149–159. https://doi.org/10.1007/s00068-006-6041-1.

[3]	W. Abd-Elaziem, M.A. Darwish, A. Hamada, W.M. Daoush, Titanium-Based alloys and composites for orthopedic implants Applications: A comprehensive review, Mater. Des. 241 (2024) 112850. https://doi.org/10.1016/j.matdes.2024.112850.

[4]	G. Mani, D. Porter, S. Collins, T. Schatz, A. Ornberg, R. Shulfer, A review on manufacturing processes of cobalt‐chromium alloy implants and its impact on corrosion resistance and biocompatibility, J. Biomed. Mater. Res. B Appl. Biomater. 112 (2024) e35431. https://doi.org/10.1002/jbm.b.35431.

[5]	M. Niinomi, M. Nakai, J. Hieda, Development of new metallic alloys for biomedical applications, Acta Biomater. 8 (2012) 3888–3903. https://doi.org/10.1016/j.actbio.2012.06.037.

[6]	M. Kaveh, M. Badrossamay, E. Foroozmehr, M. Kharaziha, Biodegradability, biocompatibility, and mechanical behavior of additively manufactured zinc scaffolds, J. Mech. Behav. Biomed. Mater. 163 (2025) 106868. https://doi.org/10.1016/j.jmbbm.2024.106868.

[7]	S. Limmahakhun, A. Oloyede, K. Sitthiseripratip, Y. Xiao, C. Yan, Stiffness and strength tailoring of cobalt chromium graded cellular structures for stress-shielding reduction, Mater. Des. 114 (2017) 633–641. https://doi.org/10.1016/j.matdes.2016.11.090.

[8]	X.P. Tan, Y.J. Tan, C.S.L. Chow, S.B. Tor, W.Y. Yeong, Metallic powder-bed based 3D printing of cellular scaffolds for orthopaedic implants: A state-of-the-art review on manufacturing, topological design, mechanical properties and biocompatibility, Mater. Sci. Eng. C 76 (2017) 1328–1343. https://doi.org/10.1016/j.msec.2017.02.094.

[9]	P. Erne, M. Schier, T.J. Resink, The Road to Bioabsorbable Stents: Reaching Clinical Reality?, Cardiovasc. Intervent. Radiol. 29 (2006) 11–16. https://doi.org/10.1007/s00270-004-0341-9.

[10]	R. Erbel, C. Di Mario, J. Bartunek, J. Bonnier, B. de Bruyne, F.R. Eberli, P. Erne, M. Haude, B. Heublein, M. Horrigan, C. Ilsley, D. Böse, J. Koolen, T.F. Lüscher, N. Weissman, R. Waksman, Temporary scaffolding of coronary arteries with bioabsorbable magnesium stents: a prospective, non-randomised multicentre trial, The Lancet 369 (2007) 1869–1875. https://doi.org/10.1016/S0140-6736(07)60853-8.

[11]	M. Hussain, S.M. Khan, M. Shafiq, N. Abbas, U. Sajjad, K. Hamid, Advances in biodegradable materials: Degradation mechanisms, mechanical properties, and biocompatibility for orthopedic applications, Heliyon 10 (2024) e32713. https://doi.org/10.1016/j.heliyon.2024.e32713.

[12]	N.S. Manam, W.S.W. Harun, D.N.A. Shri, S.A.C. Ghani, T. Kurniawan, M.H. Ismail, M.H.I. Ibrahim, Study of corrosion in biocompatible metals for implants: A review, J. Alloys Compd. 701 (2017) 698–715. https://doi.org/10.1016/j.jallcom.2017.01.196.

[13]	J. Venezuela, M.S. Dargusch, The influence of alloying and fabrication techniques on the mechanical properties, biodegradability and biocompatibility of zinc: A comprehensive review, Acta Biomater. 87 (2019) 1–40. https://doi.org/10.1016/j.actbio.2019.01.035.

[14]	L. Tan, X. Yu, P. Wan, K. Yang, Biodegradable Materials for Bone Repairs: A Review, J. Mater. Sci. Technol. 29 (2013) 503–513. https://doi.org/10.1016/j.jmst.2013.03.002.

[15]	S. Zhao, M. Tayyebi, Mahdireza Yarigarravesh, G. Hu, A review of magnesium corrosion in bio-applications: mechanism, classification, modeling, in-vitro, and in-vivo experimental testing, and tailoring Mg corrosion rate, J. Mater. Sci. 58 (2023) 12158–12181. https://doi.org/10.1007/s10853-023-08782-z.

[16]	Y. Li, H. Jahr, K. Lietaert, P. Pavanram, A. Yilmaz, L.I. Fockaert, M.A. Leeflang, B. Pouran, Y. Gonzalez-Garcia, H. Weinans, J.M.C. Mol, J. Zhou, A.A. Zadpoor, Additively manufactured biodegradable porous iron, Acta Biomater. 77 (2018) 380–393. https://doi.org/10.1016/j.actbio.2018.07.011.

[17]	J. Rao, H. Gao, J. Sun, R. Yu, D. Zhao, Y. Ding, A Critical Review of Biodegradable Zinc Alloys toward Clinical Applications, ACS Biomater. Sci. Eng. 10 (2024) 5454–5473. https://doi.org/10.1021/acsbiomaterials.4c00210.

[18]	G. Katarivas Levy, J. Goldman, E. Aghion, The Prospects of Zinc as a Structural Material for Biodegradable Implants—A Review Paper, Metals 7 (2017) 402. https://doi.org/10.3390/met7100402.

[19]	H. Kabir, K. Munir, C. Wen, Y. Li, Recent research and progress of biodegradable zinc alloys and composites for biomedical applications: Biomechanical and biocorrosion perspectives, Bioact. Mater. 6 (2021) 836–879. https://doi.org/10.1016/j.bioactmat.2020.09.013.

[20]	Y. Li, J. Zhou, P. Pavanram, M.A. Leeflang, L.I. Fockaert, B. Pouran, N. Tümer, K.-U. Schröder, J.M.C. Mol, H. Weinans, H. Jahr, A.A. Zadpoor, Additively manufactured biodegradable porous magnesium, Acta Biomater. 67 (2018) 378–392. https://doi.org/10.1016/j.actbio.2017.12.008.

[21]	N.T. Kirkland, N. Birbilis, M.P. Staiger, Assessing the corrosion of biodegradable magnesium implants: A critical review of current methodologies and their limitations, Acta Biomater. 8 (2012) 925–936. https://doi.org/10.1016/j.actbio.2011.11.014.

[22]	Y.F. Zheng, X.N. Gu, F. Witte, Biodegradable metals, Mater. Sci. Eng. R Rep. 77 (2014) 1–34. https://doi.org/10.1016/j.mser.2014.01.001.

[23]	M. Prakasam, J. Locs, K. Salma-Ancane, D. Loca, A. Largeteau, L. Berzina-Cimdina, Biodegradable Materials and Metallic Implants—A Review, J. Funct. Biomater. 8 (2017) 44. https://doi.org/10.3390/jfb8040044.

[24]	A. Joshi, G. Dias, M.P. Staiger, In silico modelling of the corrosion of biodegradable magnesium-based biomaterials: modelling approaches, validation and future perspectives, Biomater. Transl. 2 (2021) 257–271. https://doi.org/10.12336/biomatertransl.2021.03.008.

[25]	J.A. Grogan, B.J. O’Brien, S.B. Leen, P.E. McHugh, A corrosion model for bioabsorbable metallic stents, Acta Biomater. 7 (2011) 3523–3533. https://doi.org/10.1016/j.actbio.2011.05.032.

[26]	D. Gastaldi, V. Sassi, L. Petrini, M. Vedani, S. Trasatti, F. Migliavacca, Continuum damage model for bioresorbable magnesium alloy devices — Application to coronary stents, J. Mech. Behav. Biomed. Mater. 4 (2011) 352–365. https://doi.org/10.1016/j.jmbbm.2010.11.003.

[27]	F. Saconi, G.H. Diaz, A.C. Vieira, M.L. Ribeiro, Experimental Characterization and Numerical Modeling of the Corrosion Effect on the Mechanical Properties of the Biodegradable Magnesium Alloy WE43 for Orthopedic Applications, Materials 15 (2022) 7164. https://doi.org/10.3390/ma15207164.

[28]	K. Peng, X. Cui, A. Qiao, Y. Mu, Mechanical analysis of a novel biodegradable zinc alloy stent based on a degradation model, Biomed. Eng. OnLine 18 (2019) 39. https://doi.org/10.1186/s12938-019-0661-2.

[29]	Y. Li, H. Jahr, J. Zhou, A.A. Zadpoor, Additively manufactured biodegradable porous metals, Acta Biomater. 115 (2020) 29–50. https://doi.org/10.1016/j.actbio.2020.08.018.

[30]	E.B. Montufar, Bone Response to Biodegradable Metals and In Vitro Evaluation of the Cytocompatibility, JOM 77 (2025) 4473–4492. https://doi.org/10.1007/s11837-025-07353-8.

[31]	J. Zhang, Y. Jiang, Z. Shang, B. Zhao, M. Jiao, W. Liu, M. Cheng, B. Zhai, Y. Guo, B. Liu, X. Shi, B. Ma, Biodegradable metals for bone defect repair: A systematic review and meta-analysis based on animal studies, Bioact. Mater. 6 (2021) 4027–4052. https://doi.org/10.1016/j.bioactmat.2021.03.035.

[32]	Z. Wang, B. Liu, B. Yin, Y. Zheng, Y. Tian, P. Wen, Comprehensive review of additively manufactured biodegradable magnesium implants for repairing bone defects from biomechanical and biodegradable perspectives, Front. Chem. 10 (2022) 1066103. https://doi.org/10.3389/fchem.2022.1066103.

[33]	N. Pohl, D. Priebe, T. AlBaraghtheh, S. Schimek, D.C.F. Wieland, D. Krüger, S. Trostorff, R. Willumeit-Römer, R. Köhl, B. Zeller-Plumhoff, Computational modelling of bone growth and mineralization surrounding biodegradable Mg-based and permanent Ti implants, Comput. Struct. Biotechnol. J. 27 (2025) 1394–1406. https://doi.org/10.1016/j.csbj.2025.02.027.

[34]	H. Mndlovu, P. Kumar, L.C. Du Toit, Y.E. Choonara, A review of biomaterial degradation assessment approaches employed in the biomedical field, Npj Mater. Degrad. 8 (2024) 66. https://doi.org/10.1038/s41529-024-00487-1.

[35]	C. Panaghie, G. Zegan, A. Sodor, N. Cimpoeșu, N.-M. Lohan, B. Istrate, A.-M. Roman, N. Ioanid, Analysis of Degradation Products of Biodegradable ZnMgY Alloy, Materials 16 (2023) 3092. https://doi.org/10.3390/ma16083092.

[36]	T. AlBaraghtheh, R. Willumeit‐Römer, B. Zeller‐Plumhoff, Best Practices in Developing a Workflow for Uncertainty Quantification for Modeling the Biodegradation of Mg‐Based Implants, Adv. Sci. 11 (2024) 2403543. https://doi.org/10.1002/advs.202403543.

[37]	A. Amerinatanzi, R. Mehrabi, H. Ibrahim, A. Dehghan, N. Shayesteh Moghaddam, M. Elahinia, Predicting the Biodegradation of Magnesium Alloy Implants: Modeling, Parameter Identification, and Validation, Bioengineering 5 (2018) 105. https://doi.org/10.3390/bioengineering5040105.

[38]	M. Barzegari, D. Mei, S.V. Lamaka, L. Geris, Computational modeling of degradation process of biodegradable magnesium biomaterials, Corros. Sci. 190 (2021) 109674. https://doi.org/10.1016/j.corsci.2021.109674.

[39]	G. Chandra, A. Pandey, Design approaches and challenges for biodegradable bone implants: a review, Expert Rev. Med. Devices 18 (2021) 629–647. https://doi.org/10.1080/17434440.2021.1935875.

[40]	A. Cherqaoui, C. Paternoster, D. Mantovani, Biodegradable Metals and Corrosion Control: Challenges, Limits and New Opportunities for Innovating in Orthopedic Fixations, Materials 19 (2026) 1789. https://doi.org/10.3390/ma19091789.

[41]	P.K. Bowen, E.R. Shearier, S. Zhao, R.J. Guillory, F. Zhao, J. Goldman, J.W. Drelich, Biodegradable Metals for Cardiovascular Stents: from Clinical Concerns to Recent Zn‐Alloys, Adv. Healthc. Mater. 5 (2016) 1121–1140. https://doi.org/10.1002/adhm.201501019.

[42]	Z. Sheikh, S. Najeeb, Z. Khurshid, V. Verma, H. Rashid, M. Glogauer, Biodegradable Materials for Bone Repair and Tissue Engineering Applications, Materials 8 (2015) 5744–5794. https://doi.org/10.3390/ma8095273.

[43]	M. Asgari, R. Hang, C. Wang, Z. Yu, Z. Li, Y. Xiao, Biodegradable Metallic Wires in Dental and Orthopedic Applications: A Review, Metals 8 (2018) 212. https://doi.org/10.3390/met8040212.

[44]	Z. Wang, X. Wang, J. Pei, Y. Tian, J. Zhang, C. Jiang, J. Huang, Z. Pang, Y. Cao, X. Wang, S. An, X. Wang, H. Huang, G. Yuan, Z. Yan, Degradation and osteogenic induction of a SrHPO4-coated Mg–Nd–Zn–Zr alloy intramedullary nail in a rat femoral shaft fracture model, Biomaterials 247 (2020) 119962. https://doi.org/10.1016/j.biomaterials.2020.119962.

[45]	Y. Liu, T. Du, A. Qiao, Y. Mu, H. Yang, Zinc-Based Biodegradable Materials for Orthopaedic Internal Fixation, J. Funct. Biomater. 13 (2022) 164. https://doi.org/10.3390/jfb13040164.

[46]	R. Gorejová, L. Haverová, R. Oriňaková, A. Oriňak, M. Oriňak, Recent advancements in Fe-based biodegradable materials for bone repair, J. Mater. Sci. 54 (2019) 1913–1947. https://doi.org/10.1007/s10853-018-3011-z.

[47]	A.K. Nasution, M.F. Ulum, M.R.A. Kadir, H. Hermawan, Mechanical and corrosion properties of partially degradable bone screws made of pure iron and stainless steel 316L by friction welding, Sci. China Mater. 61 (2018) 593–606. https://doi.org/10.1007/s40843-017-9057-3.

[48]	N. Siroros, R. Merfort, Y. Liu, M. Praster, F. Migliorini, N. Maffulli, R. Michalik, F. Hildebrand, J. Eschweiler, Mechanical properties of a bioabsorbable magnesium interference screw for anterior cruciate ligament reconstruction in various testing bone materials, Sci. Rep. 13 (2023) 12342. https://doi.org/10.1038/s41598-023-39513-8.

[49]	M. Thomann, Ch. Krause, D. Bormann, N. von der Höh, H. Windhagen, A. Meyer‐Lindenberg, Comparison of the resorbable magnesium . alloys LAE442 und MgCa0.8 concerning their mechanical properties, their progress of degradation and the bone‐implant‐contact after 12 months implantation duration in a rabbit model, Mater. Werkst. 40 (2009) 82–87. https://doi.org/10.1002/mawe.200800412.

[50]	T. POHJONEN, P. HELEVIRTA, P. TORMALA, K. KOSKIKARE, H. PATIALA, P. ROKKANEN, Strength retention of self-reinforced poly-L-lactide screws. A comparison of compression moulded and machine cut screws, J. Mater. Sci. Mater. Med. 8 (1997) 311–320. https://doi.org/10.1023/A:1018516529591.

[51]	H. Waizy, J.-M. Seitz, J. Reifenrath, A. Weizbauer, F.-W. Bach, A. Meyer-Lindenberg, B. Denkena, H. Windhagen, Biodegradable magnesium implants for orthopedic applications, J. Mater. Sci. 48 (2013) 39–50. https://doi.org/10.1007/s10853-012-6572-2.

[52]	A. Vautrin, M. Wesseling, R. Wirix-Speetjens, M.J. Gomez-Benito, Time-dependent in silico modelling of orthognathic surgery to support the design of biodegradable bone plates, J. Mech. Behav. Biomed. Mater. 121 (2021) 104641. https://doi.org/10.1016/j.jmbbm.2021.104641.

[53]	D.H.K. Chow, J. Wang, P. Wan, L. Zheng, M.T.Y. Ong, L. Huang, W. Tong, L. Tan, K. Yang, L. Qin, Biodegradable magnesium pins enhanced the healing of transverse patellar fracture in rabbits, Bioact. Mater. 6 (2021) 4176–4185. https://doi.org/10.1016/j.bioactmat.2021.03.044.

[54]	N. Angrisani, J. Reifenrath, F. Zimmermann, R. Eifler, A. Meyer-Lindenberg, K. Vano-Herrera, C. Vogt, Biocompatibility and degradation of LAE442-based magnesium alloys after implantation of up to 3.5 years in a rabbit model, Acta Biomater. 44 (2016) 355–365. https://doi.org/10.1016/j.actbio.2016.08.002.

[55]	Y. Koo, H.-B. Lee, Z. Dong, R. Kotoka, J. Sankar, N. Huang, Y. Yun, The Effects of Static and Dynamic Loading on Biodegradable Magnesium Pins In Vitro and In Vivo, Sci. Rep. 7 (2017). https://doi.org/10.1038/s41598-017-14836-5.

[56]	S. Amukarimi, M. Mozafari, Biodegradable magnesium‐based biomaterials: An overview of challenges and opportunities, MedComm 2 (2021) 123–144. https://doi.org/10.1002/mco2.59.

[57]	V. Karageorgiou, D. Kaplan, Porosity of 3D biomaterial scaffolds and osteogenesis, Biomaterials 26 (2005) 5474–5491. https://doi.org/10.1016/j.biomaterials.2005.02.002.

[58]	H. Mohammadi, M. Sepantafar, N. Muhamad, A. Bakar Sulong, How Does Scaffold Porosity Conduct Bone Tissue Regeneration?, Adv. Eng. Mater. 23 (2021) 2100463. https://doi.org/10.1002/adem.202100463.

[59]	N. Abbasi, S. Hamlet, R.M. Love, N.-T. Nguyen, Porous scaffolds for bone regeneration, J. Sci. Adv. Mater. Devices 5 (2020) 1–9. https://doi.org/10.1016/j.jsamd.2020.01.007.

[60]	A. Motaharinia, J.W. Drelich, S. Sharif, A.F. Ismail, F. Naeimi, A. Glover, M. Ebrahiminejad, H.R. Bakhsheshi-Rad, Overview of porous magnesium-based scaffolds: development, properties and biomedical applications, Mater. Futur. 4 (2025) 012401. https://doi.org/10.1088/2752-5724/ad9493.

[61]	S. Arabnejad, B. Johnston, M. Tanzer, D. Pasini, Fully porous 3D printed titanium femoral stem to reduce stress-shielding following total hip arthroplasty: FULLY POROUS 3D PRINTED TITANIUM FEMORAL STEM, J. Orthop. Res. 35 (2017) 1774–1783. https://doi.org/10.1002/jor.23445.

[62]	W. Song, D. Zhao, F. Guo, J. Wang, Y. Wang, X. Wang, Z. Han, W. Fan, Y. Liu, Z. Xu, L. Chen, Additive manufacturing of degradable metallic scaffolds for material-structure-driven diabetic maxillofacial bone regeneration, Bioact. Mater. 36 (2024) 413–426. https://doi.org/10.1016/j.bioactmat.2024.06.028.

[63]	Y. Shi, W. Xu, H. Che, S. Zhao, W. Chang, X. Li, Y. Lu, C. Xue, D. Zhang, L.-N. Wang, Y. Li, The effect of topological design on the degradation behavior of additively manufactured porous zinc alloy, Npj Mater. Degrad. 8 (2024) 42. https://doi.org/10.1038/s41529-024-00451-z.

[64]	M. Salama, M.F. Vaz, R. Colaço, C. Santos, M. Carmezim, Biodegradable Iron and Porous Iron: Mechanical Properties, Degradation Behaviour, Manufacturing Routes and Biomedical Applications, J. Funct. Biomater. 13 (2022) 72. https://doi.org/10.3390/jfb13020072.

[65]	Y. Li, H. Jahr, P. Pavanram, F.S.L. Bobbert, U. Puggi, X.-Y. Zhang, B. Pouran, M.A. Leeflang, H. Weinans, J. Zhou, A.A. Zadpoor, Additively manufactured functionally graded biodegradable porous iron, Acta Biomater. 96 (2019) 646–661. https://doi.org/10.1016/j.actbio.2019.07.013.

[66]	J. Henkel, M.A. Woodruff, D.R. Epari, R. Steck, V. Glatt, I.C. Dickinson, P.F.M. Choong, M.A. Schuetz, D.W. Hutmacher, Bone Regeneration Based on Tissue Engineering Conceptions — A 21st Century Perspective, Bone Res. 1 (2013) 216–248. https://doi.org/10.4248/BR201303002.

[67]	C. Li, C. Guo, V. Fitzpatrick, A. Ibrahim, M.J. Zwierstra, P. Hanna, A. Lechtig, A. Nazarian, S.J. Lin, D.L. Kaplan, Design of biodegradable, implantable devices towards clinical translation, Nat. Rev. Mater. 5 (2019) 61–81. https://doi.org/10.1038/s41578-019-0150-z.

[68]	Y. Liu, B. Lu, Z. Cai, Recent Progress on Mg- and Zn-Based Alloys for Biodegradable Vascular Stent Applications, J. Nanomater. 2019 (2019) 1–16. https://doi.org/10.1155/2019/1310792.

[69]	C. Li, Y. Xian, Y. Zhou, X. Ding, J. Sun, D. Xue, Generative Design for Alloys: Harnessing Generative Models for Faster Discovery, Adv. Mater. 38 (2026) e20478. https://doi.org/10.1002/adma.202520478.

[70]	L. Sun, Y. Zeng, Z. Shen, C. Yue, Y. Yang, J. Gao, J. Zhang, Q. Yuan, L. Cha, Biodegradable Metal-Based Stents: Advances, Challenges, and Prospects, J. Funct. Biomater. 16 (2025) 315. https://doi.org/10.3390/jfb16090315.

[71]	H.Y. Ang, Y.Y. Huang, S.T. Lim, P. Wong, M. Joner, N. Foin, Mechanical behavior of polymer-based vs. metallic-based bioresorbable stents, J. Thorac. Dis. 9 (2017) S923–S934. https://doi.org/10.21037/jtd.2017.06.30.

[72]	M. Moravej, D. Mantovani, Biodegradable Metals for Cardiovascular Stent Application: Interests and New Opportunities, Int. J. Mol. Sci. 12 (2011) 4250–4270. https://doi.org/10.3390/ijms12074250.

[73]	P. Ansoms, M. Barzegari, J. Vander Sloten, L. Geris, Coupling biomechanical models of implants with biodegradation models: A case study for biodegradable mandibular bone fixation plates, J. Mech. Behav. Biomed. Mater. 147 (2023) 106120. https://doi.org/10.1016/j.jmbbm.2023.106120.

[74]	X. Yang, Z. Sun, C. Mi, Y. Hu, Optimizing magnesium scaffolds via computational modeling of degradation and bone regeneration, Mater. Today Commun. 48 (2025) 113601. https://doi.org/10.1016/j.mtcomm.2025.113601.

[75]	J. Niu, H. Huang, J. Pei, Z. Jin, S. Guan, G. Yuan, Research and development strategy for biodegradable magnesium-based vascular stents: a review, Biomater. Transl. 2 (2021) 236–247. https://doi.org/10.12336/biomatertransl.2021.03.06.

[76]	G.S. Karanasiou, M.I. Papafaklis, C. Conway, L.K. Michalis, R. Tzafriri, E.R. Edelman, D.I. Fotiadis, Stents: Biomechanics, Biomaterials, and Insights from Computational Modeling, Ann. Biomed. Eng. 45 (2017) 853–872. https://doi.org/10.1007/s10439-017-1806-8.

[77]	K.K. Alaneme, S.A. Kareem, J.L. Olajide, R.E. Sadiku, M.O. Bodunrin, Computational biomechanical and biodegradation integrity assessment of Mg-based biomedical devices for cardiovascular and orthopedic applications: A review, Int. J. Lightweight Mater. Manuf. 5 (2022) 251–266. https://doi.org/10.1016/j.ijlmm.2022.02.003.

[78]	A. Kansal, N. Dixit, S. Kumar, A. Dvivedi, P. Kuamr, Review: zinc alloys for biomedical applications—alloy design, processing techniques, and bio-functional performance, J. Mater. Sci. 61 (2026) 3676–3713. https://doi.org/10.1007/s10853-025-12110-y.

[79]	C. Shuai, S. Li, S. Peng, P. Feng, Y. Lai, C. Gao, Biodegradable metallic bone implants, Mater. Chem. Front. 3 (2019) 544–562. https://doi.org/10.1039/C8QM00507A.

[80]	V. Zarghami, Synchronizing degradation with regeneration: a model-driven framework for designing biodegradable biomaterials in bone tissue engineering, J. Mater. Sci. Mater. Med. (2026). https://doi.org/10.1007/s10856-026-07102-6.

[81]	Y. Su, J. Fu, W. Lee, S. Du, Y.-X. Qin, Y. Zheng, Y. Wang, D. Zhu, Improved mechanical, degradation, and biological performances of Zn–Fe alloys as bioresorbable implants, Bioact. Mater. 17 (2022) 334–343. https://doi.org/10.1016/j.bioactmat.2021.12.030.

[82]	M. Krämer, M. Schilling, R. Eifler, B. Hering, J. Reifenrath, S. Besdo, H. Windhagen, E. Willbold, A. Weizbauer, Corrosion behavior, biocompatibility and biomechanical stability of a prototype magnesium-based biodegradable intramedullary nailing system, Mater. Sci. Eng. C 59 (2016) 129–135. https://doi.org/10.1016/j.msec.2015.10.006.

[83]	Z. Shen, M. Zhao, X. Zhou, H. Yang, J. Liu, H. Guo, Y. Zheng, J.-A. Yang, A numerical corrosion-fatigue model for biodegradable Mg alloy stents, Acta Biomater. 97 (2019) 671–680. https://doi.org/10.1016/j.actbio.2019.08.004.

[84]	J. Qi, H. Zhang, S. Chen, T. Du, Y. Zhang, A. Qiao, Numerical Simulation of Dynamic Degradation and Fatigue Damage of Degradable Zinc Alloy Stents, J. Funct. Biomater. 14 (2023) 547. https://doi.org/10.3390/jfb14110547.

[85]	S. Söntgen, L. Keilig, K. Kabir, A. Weber, S. Reimann, K. Welle, C. Bourauel, Mechanical and numerical investigations of biodegradable magnesium alloy screws for fracture treatment, J. Biomed. Mater. Res. B Appl. Biomater. 111 (2023) 7–15. https://doi.org/10.1002/jbm.b.35127.

[86]	S. Byun, H. Lim, K. Cheon, S. Lee, H. Kim, J. Lee, Biodegradable magnesium alloy (WE43) in bone‐fixation plate and screw, J. Biomed. Mater. Res. B Appl. Biomater. 108 (2020) 2505–2512. https://doi.org/10.1002/jbm.b.34582.

[87]	J. Tao, J. Cheng, G. Zhao, X. Hu, X. Gu, J. Su, Computational Modeling of the Corrosion Process and Mechanical Performance of Biodegradable Stent, J. Phys. Conf. Ser. 1888 (2021) 012019. https://doi.org/10.1088/1742-6596/1888/1/012019.

[88]	M. Peuster, C. Hesse, T. Schloo, C. Fink, P. Beerbaum, C. von Schnakenburg, Long-term biocompatibility of a corrodible peripheral iron stent in the porcine descending aorta, Biomaterials 27 (2006) 4955–4962. https://doi.org/10.1016/j.biomaterials.2006.05.029.

[89]	Y. Jin, J. Li, H. Fan, J. Du, Y. He, Biomechanics and Mechanobiology of Additively Manufactured Porous Load‐Bearing Bone Implants, Small 21 (2025) 2409955. https://doi.org/10.1002/smll.202409955.

[90]	K. Chen, X. Gu, Y. Zheng, Feasibility, challenges and future prospects of biodegradable zinc alloys as orthopedic internal fixation implants, Smart Mater. Manuf. 2 (2024) 100042. https://doi.org/10.1016/j.smmf.2023.100042.

[91]	Z. Wang, Z. Lv, X. Cai, Y. Wang, B. Peng, H. Xu, H. Pang, X. Yang, J. Xu, Y. Bian, B. Feng, P. Wen, Y. Zheng, X. Weng, Sculpting the Future of Bone: The Evolution of Absorbable Materials in Orthopedics, Adv. Mater. 38 (2026). https://doi.org/10.1002/adma.202510848.